- Purpose

This procedure provides instruction for the development and completion of software validation.

- Scope

Software validation is required to validate new or significantly modified software that is intended as a part of a finished medical device. This procedure is intended to be followed by all personnel who are involved with software development.

- Definitions

| Design Validation | Establishing by objective evidence that device specifications conform to user needs and intended uses. |
| --- | --- |
| Process Validation | Establishing by objective evidence that a process consistently produces a result or product meeting its predetermined specifications. |
| Requirements | The needs or expectations for a system and/or its software. There can be many different kinds of requirements (e.g., design, functional, implementation, interface, performance or physical requirements). |
| Specification | Any requirement with which a product, process, service or other activity must conform. |
| Verification | Confirmation by examination and provision of objective evidence that specified requirements have been fulfilled. |
| Validation | Confirmation by examination and provision of objective evidence that the particular requirements for a specific intended use can be consistently fulfilled. |

- Responsibilities

| User | - Complies with the policy and procedure for all new or revised design projects. - Ensures the most current version of this document is used when referenced. - Acts as project coordinator to ensure proper protocol development, accurate and timely protocol execution, and development of the final report. |
| --- | --- |
| Departmental Management | - Ensures departmental personnel are properly trained before using this policy or procedure. - Assigning personnel as project leader and/or to validation team(s). - Complies with the policy and procedure for all new or revised design projects. - Ensures the most current version of this document is used when referenced. - Reviews and approves protocols prior to validation and summary report of the validation activities to ensure that activities reflect actual activities performed, at a minimum. |
| Quality Assurance | - Complies with and enforces the document control standards within this document. - Monitors the implementation and effectiveness of this procedure. - Conducts periodic audits to assure compliance. - Reviews and approves protocols prior to validation and summary report of the validation activities to ensure that all quality system requirements are met, at a minimum. |
| Engineering/Manufacturing | - Provides technical and specific guidance for protocol development and validation process execution. - Reviews and approves protocols prior to validation and summary report of the validation activities to ensure that summary report meets the technical requirements for software validation and that conclusions drawn are valid, at a minimum. |

- Policy

When/if computers or automated data processing systems are used as part of a finished device or a part of EPIC’s quality system, the software used is validated for its intended use. New processes or changes/deviations to existing software are verified and/or validated prior to release of the products. Validation requirements are reviewed and addressed for each product to ensure adherence to all proposed uses and functions according to applicable specifications and requirements. Validation is required to be performed and documented for process or design changes except when verification testing is acceptable. Verification testing may be acceptable in such instances as minor code changes or where changes can be 100% verified. The decision to perform validation versus verification testing is made on a case-by-case basis and is determined by reviewing the specific changes and its effect on form, fit, function, safety and effectiveness. All decisions regarding validation versus verification will be documented. Approved validation and verification documents are maintained and filed in the appropriate design history file (DHF) maintained by Quality Assurance (QA).

- Procedure

  - IT completes software verification testing as outlined in EG-016, Software Verification Procedure.

  - System tests as a part of design verification are controlled via EG-016. Validation tests are controlled by this procedure. Validation test requirements are established in the verification and validation test plan.

NOTE: Clinical investigations may be used to satisfy validation requirements.  Clinical investigations are controlled by clinical staff according to the appropriate clinical procedures. As such, clinical protocols and final reports must comply with clinical procedures which may be different from those specified below.

  - Validation protocols will be developed and approved using EPIC Engineering Studies procedure.

  - Protocol approval will be obtained from all appropriate functional groups as specified in a design and development plan.

  - Quality Assurance will coordinate with IT to ensure that the test machines have a completely installed version of the software. IT will document the test equipment set-up using the standard operating procedure for software installation. This documentation will be included in the final summary report.

  - Protocol Execution

    - QA will coordinate execution of the software validation test protocols. All records will be retained and included in the final summary report.

    - QA and IT will coordinate the investigation and resolution of any non-conformances and/or deviations discovered during validation testing. If code changes are required to resolve non-conformances, IT will modify the code and issue a new software version which will then be deployed for use in retesting.

    - The Revision History Record, as defined by EG-001, Design Control Procedure, will be used to track software versions created to resolve issues discovered during validation testing.

  - Final Software Validation Test Report

    - Upon completion of the validation testing, a summary of all test plans, expected outcomes, and test results including the summary of any version changes executed during development will be documented in a summary report using EPIC’s Engineering Studies procedure.

    - A validation analysis may to be included in the final report to establish documentation to summarize multiple validation test results.

    - The final report approval will include the same functional groups included in the test protocol approval.

  - Upon approval of the final report of all validation tests, QA will forward the final reports to the project leader for inclusion in the Product Release Review, as defined by EG-001.

- Reference

EG-001, Design Control Procedure

EG-016, Software Verification Procedure

EG-020, Engineering Studies

- Attachments

N/A

Document Revision History

| Change Order Number: | Version Number: | Description of Change: | Date: | Updated by: |
| --- | --- | --- | --- | --- |
| CC-011810-01 | 000 | New Document | 1/26/2010 | AFM |
| CC-120202-01 | 001 | Rewrite procedure to remove reference to tracking in the Gemini system and reflect current design development activities. | 2/3/12 | A. Mason |
| CC-120514-02 | 002 | Rewrite procedure to reflect current practices and clearly indicate order of operations. Add reference to new Engineering Studies procedure. | 5/14/12 | A. Miller |
| CC-130319-02 | 003 | Update procedure to reflect changes made to the design control processes separating system tests as a part of design verification outlined in EG-016. | 3/21/13 | L. Kurth |