# Purpose

The purpose of this document is to define the Software Verification Procedure that will be followed for software product development at EPIC.

# Scope

This procedure is intended to be used for verification of the software components of EPIC products. This procedure applies to the design and development of any:

Software used as a component, part, or accessory of a medical device.

Software that is itself a medical device

Software used in the production of a device; and

Software used in the implementation of the device manufacturer’s quality system

Software intended to support any activity as a part of clinical (IDE) data collection

This procedure is intended to be followed by all personnel who are involved with software design and development projects.

# Definitions

| Design History File (DHF) | A compilation of records which describes the design history of a finished device. |
| --- | --- |
| Design Phase | A series of stages, including design, implementation, verification and reviews that results in the production of a software version. See EG-001 for full definition. |
| Design Review | A documented, comprehensive, and systematic examination of a design to evaluate the adequacy of the design requirements, to evaluate the capability of the design to meet these requirements, and to identify problems. |
| Project Leader | The member of the Project Team who coordinates the work of the team and reports progress and project status to management. |
| Project Team | The set of individuals who together collaborate to develop a product. This document uses the term to describe the team developing software. See EG-001 for full definition. |
| Regression Testing | The selective retesting of a previously tested software system to ensure that changes to the software have not introduced new errors or defects. |
| Requirement | Any need or expectation for a system or for its software. Requirements reflect the stated or implied needs of the customer, and may be market-based, contractual, or statutory, as well as an organization’s internal requirements. |
| Software Validation | Confirmation by examination and provision of objective evidence that software specifications conform to user needs and intended uses, and that particular requirements implemented through software can be consistently fulfilled. |
| Software Verification | Provides objective evidence that the design outputs of a particular phase of the software development life cycle meet all the specified requirements for that phase. |
| Version | A specific, documented edition of a software application produced as an output of a Design Phase. |

# Procedure

This Software Verification Procedure is not a step-by-step procedure like the Design Control Procedure, as defined in EG-001. Software verification looks for consistency, completeness, and correctness of the software and its supporting documentation, as it is being developed, and provides support for a subsequent conclusion that software is validated.

This procedure defines the verification activities that, when performed, will lead to that conclusion. Where appropriate, it defines when each of the activities should be performed.

## Design Control Procedure

Software verification is part of the Design Control Procedure. Conversely, the Design Control Procedure is a part of verification. The Design Control Procedure provides the structure and discipline for systematic software development. This structure, along with its oversight, is a key ingredient to reaching a conclusion that the software is validated.

## Design Reviews

EG-001 defines four formal Design Reviews: Product Design Review, Phase Design Review, Phase Release Review, and Product Release Review. In addition, for software, there is a Software Design Evaluation. These Design Reviews are verification activities and provide support for a conclusion that software is validated.

## Traceability Analysis

The Traceability Analysis Document (TAD) ties together Design Inputs (primarily requirements) and Design Outputs (design documents and test designs/plans). Initially, the TAD provides input for the design and development of verification tests, particularly System Tests, as defined below, since it allows elements of the software design to be related to specific requirements. After tests have been executed, it allows the test results to be traced back through the design to specific requirements. In the end, the goal is to show the software, verified by the tests, satisfies the requirements. The TAD shows how the results of tests verify the implementation of requirements.

## Unit Tests

Unit Tests are structural tests and are also known as “white-box” tests. The term “white-box” indicates that the person defining and/or executing a test is doing so with knowledge based on the source code and other design documentation.

Unit Tests are low-level tests performed on artifacts of the implementation. In common object-oriented systems, Unit Tests would execute on individual methods of classes, or instances of a class (object) when a class can be initialized in multiple ways and its internal logic is dependent on its initialization. The methods and classes are the “units.”

The goal of Unit Tests is to exercise each unit of software, providing varying inputs and checking for correct outputs, to achieve a high degree of confidence that all possible logic paths have been tested. In reality, it is not possible to test every potential path with every potential input. Thus, Unit Tests should focus on the following areas:

### Statement Coverage

There should be sufficient test cases for each program statement to be executed at least once.

### Decision (Branch) Coverage

There should be sufficient test cases for each program decision outcome, or branch, to be executed at least once so that each possible outcome occurs at least once.

### Condition Coverage

There should be sufficient test cases for each condition in a program decision to take on all possible outcomes at least once. It differs from branch coverage only when multiple conditions must be evaluated to reach a decision. For example, given the compound condition “if A is true or B is true,” there are two possible outcomes: the compound condition is true or the compound condition is false. So a test case with A true and B false (or true) yields the former and a test case with A false and B false yields the latter.

### Multi-Condition Coverage

There should be sufficient test cases to exercise all possible combinations of conditions in a program decision. For example, given the compound condition “if A is true or B is true,” there are four combinations of A and B: both true, both false, A true and B false, A false and B true.

### Loop Coverage

There should be sufficient test cases for all program loops to be executed for zero, one, two and many iterations covering initialization, typical running and termination conditions.

### Path Coverage

There should be sufficient test cases for each feasible path, from start to exit of a defined unit, to be executed at least once. Full (100%) path coverage is generally not achievable because of the very large number of paths through a software system. The amount of path coverage required should be established based on risk and criticality of the unit under test.

### Data Flow Coverage

There should be sufficient test cases for each feasible data flow to be executed at least once.

In general, larger units will have a higher degree of complexity making them more difficult to test with good coverage. As complexity grows the cost of testing, in terms of time and resources, grows exponentially. Thus, Unit Tests are constructed with a bottom-up approach.

Particular focus should be given to the smallest of units, e.g. methods that are self-contained and do not require invocation of other methods. It should be possible to define test cases for these units that will provide full coverage. Moving up a level, test cases for other methods that invoke the most primitive methods should cover the possible outputs for each of the primitive methods. Unit Tests for each successive higher level will become focused on coverage for each unit at that level, not the units at lower levels. In this way, confidence is built knowing that coverage at each lower level is sufficient that no unexpected outputs will be encountered.

Unit Tests for larger functional units may be considered integration tests. These tests, exercising functional units, will test how the functional units interact with each other.

Unit Tests are normally designed and written as software code. A module of Unit Test code invokes the methods/functions of the code under test. Often Unit Tests are constructed using a Unit Test framework that contains the functionality to start and stop packages of Unit Tests and record results in logs.

Designing and writing of Unit Tests is part of the Software Construction stage defined by EG-001. Unit Tests for software code will be available for review along with the code at Code Reviews.

Unit Test execution may be automated as much as possible so that Unit Tests can be used as regression tests when code changes. Ideally, the full suite of Unit Tests will execute whenever the software system is full compiled, or built. Unit Tests can be initiated after a successful build and allowed to run and record results to a log file that can then be examined for any failures.

## Code Reviews

A Code Review, also known as a code inspection or code walkthrough, is a systematic examination of the software code to assess its conformance to standards, as defined in EG-022, and to detect errors. Code Reviews will be conducted at the end of Software Construction when code changes and Unit Tests are complete.

### The review team for a Code Review will be the developer who wrote the code (author) and at least one other independent developer, preferably more than one.

### The review will consist of two activities: 1) each reviewer will individually review the code under review along with associated unit tests; they may have discussions with each other during this activity, and 2) a formal meeting attended by all reviewers to document findings.

### Prior to the review, the author will notify review team members of the code modules to be reviewed. The notification may be electronic; developers may review code using development tools instead of hard copy.

### Each individual reviewer will review the code and Unit Tests by the following criteria; all criteria must be met for the code to pass review and be approved:

Does the code, as written, conform to established standards?

Is the functionality implemented by the code correct according to design specifications?

Are the Unit Tests for the code sufficient to provide a high degree of confidence that all possible logic paths have been tested?

### The agenda for Code Review meetings will be the following:

Discussion of code issues, if any, found during review

Discussion of unit test issues, if any, found during review

Assignment of action items to correct issues, if any

Approval summary

### Meeting minutes will document all the agenda items. EGF-016, Software Code Review Form should be used for this documentation.

### Issues that require code corrections will require the corrections to be reviewed. This same process applies with the following exceptions:

EG-016, or similar, will indicate the review is for corrections, not an initial review

Code changed for issue resolution must be reviewed.  Changes will be evaluated to determine if any unintentional effect on other code is possible. If other code may be affected, it will also be reviewed.

Justification for limiting the code review to only the changes, or extending it to include other code, will be documented.

### Documentation from the review meeting will be placed in the DHF.

The following guidelines should be used when determining the amount of code that should be subjected to review during one Code Review:

An individual reviewer should be able to review the code under review in no more than 3 hours

Large amounts of code should be separated into logical review packages and each package reviewed separately

Code under review should be related functionally

Code for multiple minor changes may be grouped together even though it isn’t related

## System Tests

System Tests are functional tests and are also known as “black-box” tests. The term “black-box” indicates that the person writing test cases and/or performing tests is treating the system as a “black,” or opaque box. Test cases are written with the knowledge of requirements and design specifications that describe the design of the software, but with no knowledge of the internals or implementation.

System Tests are high-level tests performed on the system as a whole. System Tests exercise the system in its operating environment and challenge its intended use (functionality). It isn’t possible to test every potential logic path with every potential input. System Tests should focus on the following areas:

### Normal Case

There should be sufficient test cases to cover all usual inputs and usual combinations of inputs.

### Output Forcing

There should be sufficient test cases to ensure that most software outputs are generated. It may not be feasible to generate all outputs.

### Robustness

There should be sufficient test cases to ensure the software behaves correctly when given unexpected, invalid inputs. Focus should be given to boundary values, where the values are on the boundaries of acceptable inputs, as well as inputs that are well beyond the range of acceptable input.

### Combination of Inputs

There should be sufficient test cases to exercise the software with combinations of inputs. The combinations should include those that are related and those that may be unrelated to one another.

Designing and writing of System Tests should be done during the Software Design stage (defined by EG-001), or in parallel to the Software Construction stage, using only the requirements and software design specifications as input. Ideally, System Tests will be designed and written by test engineers who are solely focused on testing and are independent from the software developers who design and implement the software code. However, in small organizations this may not be possible. In such cases, if the software developers are also writing the system tests, care must be taken to write them early, and based on the requirements and design, not the detailed knowledge gained from software construction. System Test plans/specifications should be formally reviewed by independent reviewers.

System Tests will be considered Verification Protocols and will be approved at the Phase Design Review. They will be executed, and results documented, during the Verification stage. (See EG-001 for definitions of Verification Protocols, Phase Design Review, and the Verification stage.)

System test protocols and final reports will be documented as outlined in EG-020, Engineering Studies procedure.

## Undocumented Tests

Ad hoc testing occurs naturally as part of software development. For example, a developer may write some short tests to run against a code library to confirm his/her understanding of how the library functions. These tests may be undocumented. Unless specifically enumerated in this document, or another quality system document, it is not necessary to document these tests. In some cases, the project team may decide to formally document other tests for the benefit of the project and the team as a whole.

## Software Maintenance

Software maintenance consists of changes to previously verified software to correct errors and faults (corrective maintenance), to improve performance and maintainability (perfective maintenance), and to adapt the system to be usable in a different environment (adaptive maintenance). Any time a change is made to a previously verified software system, sufficient regression testing is necessary to ensure that portions of the software not involved in the change were not adversely affected by the change. This regression testing is in addition to tests that verify the change.

Tests required as a result of change implementation will be determined when assessing the impact of the change on the software design. This will normally be during the investigation identified as the first stage of a Design Phase, as defined by EG-001. It should be determined during investigation what changes are required to Unit Tests and System Tests. Also, it should be determined which tests are necessary to be executed for regression testing. These findings should be documented in the Phase Plan for the Design Phase.

## Design History

All documentation of verification activities will become part of the product’s Design History File, as defined by EG-001. Documentation should include all of the following, as appropriate:

Plans that identify verification activities

All Design Review documentation

Any generated reports from the running of Unit Tests; the report from a final run prior to the Phase Design Review is sufficient

Minutes from Code Review meetings

Approvals from Code Reviews

System Test design/plan documentation

Results/Reports from the execution of System Tests

Other documented test activities.

# References

| FDA | General Principles of Software Validation; Final Guidance for Industry and FDA Staff. January 11, 2002 |
| --- | --- |
| EG-001 | Design Control Procedure |
| EG-020 | Engineering Studies |
| EG-022 | Software Development Guidance |
| EGF-002 | Traceability Analysis |
| EGF-014 | Design Change Approval Form |
| EGF-016 | Software Code Review Form |

Document Revision History

| Change Order | Version | Description of Change | Date | Updated by: |
| --- | --- | --- | --- | --- |
| CC-120202-01 | 000 | New Document | 2/3/12 | A. Mason |
| CC-120328-01 | 001 | Update the policy section to match current verification testing expectations. Update section 6.2 to provide clarification and examples on the record keeping requirements. Remove requirements for next steps and refer user to the Software Development SOP. | 5/14/12 | A. Mason |
| CC-130319-02 | 002 | Rewritten to encompass all aspects of verification throughout the development process. Renamed to be Software Verification Procedure. | 3/21/13 | L. Kurth |