- PURPOSE

The purpose of this document is to outline a procedure for Design Control at EPIC

- SCOPE

This procedure is intended to be used to control the design of new products.  This procedure applies to the design and development of:

- Medical device(s)

- Component(s), part(s), or accessory(ies) of a medical device;

This procedure is intended to be followed by all EPIC personnel who are involved with design and development projects.

Each proposed change to an existing product or process design will be evaluated to ensure that appropriate design control elements are applied to the proposed change.

- Definitions

| Class I Change | Any change that affects form, fit or function or that is not interchangeable. This category of change applies to new products and/or changes to the function of the existing products. |
| --- | --- |
| Design Validation | Establishing by objective evidence that device specifications conform to user needs and intended uses. |
| Device Master Record (DMR) | A compilation of records of records containing the procedures and specifications for a finished device. |
| Process Validation | Establishing by objective evidence that a process consistently produces a result or product meeting its predetermined specifications. |
| Requirements | The needs or expectations for a system and/or its software. There can be many different kinds of requirements (e.g., design, functional, implementation, interface, performance or physical requirements). |
| Specification | In software development, a specification is defined as a document that states the requirements |
| Verification | Confirmation by examination and provision of objective evidence that specified requirements have been fulfilled. |
| Validation | Confirmation by examination and provision of objective evidence that the particular requirements for a specific intended use can be consistently fulfilled. Also said as, confirmation by examination and provision of objective evidence that software specifications conform to use needs and intended uses, and that the particular requirements implemented through software can be consistently fulfilled. |

- RESPONSIBILITIES AND APPROVALS

  - Design Review Board (DRB):

    - The DRB is to perform independent reviews of the new product or process or proposed changes (Class I) to existing product or process design.

    - DRB consists of a representative from (as specified in the Design and Development Plan (DDP)):

- Manufacturing / Operations

- Quality Assurance

- Regulatory Affairs

- Technical Area (e.g., Engineering)

- Independent Technical Area Review

- Other Individuals as appropriate

    - Ensures that ‘Inputs’ are established and approved.

NOTE:  The Requirements Document is to be approved by the DRB as revisions occur.  The initial approval of the Design and Development Plan and the Initial Risk Analysis is to be approved by DRB.  Subsequent changes to these two documents can be approved by the project team and/or Project Leader.

    - Ensures that ‘Outputs’ meet the approved ‘Inputs’.

    - Authorizes the design transfer into production.

    - Approves changes to the Design Control elements or other aspects that affect the design of the product.

  - Project Team:

    - Establishes, plans, coordinates and carries out the design control elements.

    - Approves changes and up-dates to schedules, plans, risk analysis and other project related changes that affect the design project.  See note on 4.1.3.  Changes and updates to schedules, plans, risk analysis will be approved by the Project Team and will be maintained by the Project Leader.  These documents will be included as part of the Design History file and / or the Final Design Review Packet.

  - Quality Assurance:

    - Maintains the completed Design History File for each new product or process design change once Design Transfer has been completed.

    - Assesses new raw material suppliers.

- DESIGN CONTROL ELEMENTS

See Design Control Flow Chart

  - Design Control Statement

  - Design Inputs:

- Requirements Document

- Initial Risk Analysis Document

- Design and Development Plan

  - Design Input Approval:

- Initial Design Review

  - Design Outputs:

- Design Verification

- Design Validation

- Final Risk Analysis Document

  - Input/output Document

  - Design Output Approval:

- Final Design Review

NOTE:  Multiple Design Reviews may be held throughout the Design Control process in addition to the Initial and Final Design Reviews.

  - Design Transfer

  - Design History File

- DEFINITIONS

  - Design Control Statement:  A summary of the Design Control elements that will be employed in the Design Control of a new product or change to an existing product or process.  If a Design Control element (i.e. Requirements Document, Risk Analysis, Design and Development Plan, Design Verification, Design Validation, Design Transfer, Design History File) is deemed as not applicable in the Design Control process, a rationale and justification is to be provided in this statement.

  - Design Inputs:  All the important and critical requirements of a product or plan that addresses the intended use of the product, including the needs of the user and patient.  The Design Input for the product includes the physical and performance requirements of a product that are used as a basis for product design. The Design Inputs may be comprised of the elements contained in the Requirement document, the initial Risk Analysis document and the Design and Development plan.

  - Requirements Document:  The Requirements document lists out all the requirements necessary for the new product or change to the product or process relating to the needs of the user and patient.  Requirements may include elements such as product functionality, product performance, safety, efficacy, storage; testing, stability, validation, regulatory, biocompatibility, handling and interfaces of the product with external sources (i.e. companion products, instruments used with the product).  This document is intended to clarify and provide specific requirements that are clear and unambiguous.

  - Initial Risk Analysis Document:  The Initial Risk Analysis provides documented evidence of an evaluation of the potential risks associated with the new product or change was performed.  The Risk Analysis identifies the failure modes that could render the product to become non-efficacious or unsafe for use.  This assessment provides a method for measuring the degree of risk associated with the new product design or change.

  - Design and Development Plan:  A plan that should consist of elements such as objectives, assumptions, major tasks, milestones, key decision points, resources, responsibilities, contract services and interrelationships between tasks and groups and identification of major design reviews.  The plan shall be reviewed, up-dated and approved as the design and development evolves.

  - Design Input Approval:  An approval of the Design Inputs by the Design Review Board.  Approval of the Design Inputs provides direction for the project team to work toward the appropriate product or process design.

  - Initial Design Review:  A Design Review is a documented, comprehensive, systematic examination of a design to evaluate the adequacy of the design requirement, to evaluate the capability of the design to meet these requirements and to identify problems.

  - An initial review and approval of the Design Inputs:  This review may include the review of the Requirements Document, Initial Risk Analysis and the Design and Development plan.  It is the Design Review Board that participates in the initial Design Reviews.

  - Design Outputs:  The result of a design effort at each design phase and at the end of the total design effort.  The total design Output consists of the product, its packaging and labeling and the device master record.All the important and critical outcomes from the development process relating to the intended use of the product or changes that could affect the intended use of the product.  The Design Outputs may be comprised of the elements contained in the Design Verification, Design Validation and the Final Risk Analysis documents.

  - Design Verification:  Design Verification is defined as a confirmation and examination that provisions of objective evidence that specified requirements have been fulfilled.  Design Verification demonstrates that the design output meets the design input requirements.  Design Verification may include the testing, experimentation, analysis and/or evaluations to demonstrate that design outputs meet the design inputs.

  - Design Validation and Process Validation:  Design Validation includes the establishment by objective evidence that the product specifications conform with user needs and intended use of the product.  Design validation also includes testing to defined operating conditions under actual or simulated product conditions.  Process Validation is defined as the establishment by objective evidence that a process consistently produces a result or product meeting its pre-determined specifications.

  - Final Risk Analysis:  The Final Risk Analysis provides documented evidence that a thorough review of the information collected throughout the Design Verification and Design Validation addresses the potential failure modes and an evaluation of the potential risks associated with the new product or change after design has taken place.

  - Input/output Document:  A summary of the design inputs and associated design outputs with references to appropriate evaluations or analysis.  This document is intended to show in a summary format that all the design inputs have been considered and met.

  - Design Output Approval:  The approval that the Design Inputs have been met as demonstrated by the Design Outputs.

  - Final Design Review:  A final review and approval that the Design Outputs meet the Design Inputs.  This review may include the review of the Final Requirements Document, Risk Analysis, the Design and Development plan, Design Verification, Design Validation and the Input/Output document.  It is the Final Design Review Board that participates in the Final Design Reviews.

  - Design Transfer:  The implementation of a new product or a change to an existing product or process into production for commercial use.  Although Design Transfer is described at the end of the Design Control, the Design Transfer process should be occurring throughout the entire development process.  Training, familiarization, discussions and feedback between the Manufacturing groups, other groups and the product designers should be occurring throughout the development process.

  - Design History File:  A compilation of records and documents that describes the design history of the finished product.  The Design History File shall contain or reference all the design control elements required for the development of a new product and or changes to an existing product or process.

  - Class I Change: Any change that affects form, fit or function or that is not interchangeable.  This category of change applies to new products and/or changes to the function of the existing products.  (Significant changes proposed after initial design is approved needs to be considered in context of change control and design control policy.)

- PROCEDURE

  - For software development, use EPIC’s Software Development procedure.  This procedure is intended for non-software development projects.  If a project involves both hardware and software development, project teams will use the guidance from this and the Software Development procedure to ensure that all appropriate design controls are exercised during the development process.

  - User/Customer ‘Essentials’ and ‘Desirables’:  An assessment of the User/Customer ‘Essentials’ and ‘Desirables’ should be carried out for a new product.  As appropriate, feedback from the customer should be considered for changes to existing products or process. Customer ‘Essentials’ and ‘Desirables’ may be defined as the requirements from patients, physician, hospitals, regulatory agencies,  sales/marketing groups, manufacturing groups (for manufacturability) or other appropriate sources. These ‘Essentials’ and ‘Desirables’ may include the product performance/physical/chemical characteristics, safety/efficacy characteristics, product functionality,  reliability, compatibility, packaging, labeling and handling.

  - Based on the feedback from the Customer as outlined above, concept identification and/or feasibility evaluations may be conducted.  Multiple product design alternatives may be considered and evaluated at this stage.

  - Once the concept and feasibility activities have been completed, the Design Control process begins.

  - Initially, an assessment should be conducted to determine which of the Design Control elements apply to the Design Control process.  All the Design Control elements will apply to a new product.  Changes to existing products or process may not require that all elements of Design Control be addressed.  Changes to an existing product are to be assessed and determined if the changes would be considered a Class I change.  This assessment is to be conducted and documented on the Design Control Statement.

  - The Design Control Statement may be included in the Design and Development Plan (DDP) or it may be a separate, standalone document.  This statement documents the Design Control elements that will be employed.  If a Design Control element (i.e. Requirements Document, Risk Analysis, and Design Validation) is deemed as not applicable in the Design Control process, a rationale and justification is to be provided in this document.  This document is to be approved and included in the Design History File.

  - A Requirements Document is to be established based on the ‘Essentials’ and ‘Desirables’ of the user/customer(s).  General information from the User/Customer ‘Essentials’ and ‘Desirables’ activity is to be reduced into detailed requirements.  These requirements need to be verifiable and validatable as appropriate. This Requirement Document is intended to provide fairly specific requirements that can include:

- Intended Use

- Description

- Product Characteristics, e.g. physical, mechanical, safety, etc.

- End-User requirements

- Product functional characteristics

- Regulatory requirements

- Manufacturing requirements

- Methods requirements

- Test/release requirements

- Packaging and labeling requirements

- Storage and handling requirements

- Shelf life requirements

- Shipping/Transportation requirements

- Servicing/replacement requirements

NOTE:  Initially, the Requirement Document may not be complete in all necessary areas.  As more information becomes available, the Requirement Document can be modified, added to and expanded.

After the initial Requirement Document is approved, the subsequent Requirement Documents must incorporate change control and be approved by the DRB.

  - An Initial Risk Analysis of the product or process is to be conducted.  The product requirements or process function (derived from the Requirement Document) is to be evaluated for potential failure modes, potential effects of a failure, potential cause of the failure and current/planned process controls. The projected severity, occurrence and detection of failure are to be evaluated.  Based on this assessment, failure risk can be assigned to product design or process aspects.  EPIC’s Risk Assessment procedure should be used to guide this risk analysis assessment.

  - A Design and Development plan is to be prepared to ensure that the design process is appropriately thought-out and controlled.

    - The plan should consist of elements such as goals/objectives, assumptions, major tasks, deliverables, milestones, key decision points, resources, responsibilities, contract services and interrelationship between tasks and groups.  The plan is to include the identification of design review throughout the project.

    - The plan is to include information on the key design control elements such as Design Inputs, Design Input approvals, Design Outputs, Design output approvals and Design transfer.

NOTE:  Initially, the Design and Development plan and initial Risk Analysis may not be complete in all necessary areas.  As more information and analysis becomes available, the plan can be modified, added to, deleted from and expanded.  After the initial Design and Development Plan and Risk Analysis is approved, substantial changes to the plan are to be reviewed and approved by the DRB.

  - When the Design Inputs have been determined (Design Inputs may include the Design Control Statement, Requirement Document, Initial Risk Analysis and the Design and Development Plan), a review is to be conducted.

    - The initial Design Review should include:

- An agenda

- An objective

- Presentation of documents, results or other information pertinent to the Design Review

- Conclusions and/or Recommendations

- Outstanding items and assignments as appropriate

- Approvals

    - A Design Review packet should be prepared in advance and distributed to the Design Review Board.  The Board should review the Design Review packet and be prepared with comments, feedback, questions and any concerns.

    - Minutes and/or actions and assignments coming from the Initial Design Review are to be documented.  Outstanding issues, if any, are to be addressed.  If outstanding actions from the Design Review cannot be closed out and addressed before the Design Review Packet is approved, these actions may be tracked as preventative actions according to CAPA procedure.

    - If there are no issues or all issues can be resolved at the meeting, the Design Review Board will sign and date signifying ‘acceptance/approval’ of the initial Design Review packet.

    - If all issues cannot be resolved at the meeting, actions will be identified and assignments made to resolve the issues.  Minor issues can be addressed and the Design Review packet can be circulated at a later time for approval.  If major issues are identified, it may be necessary to revise the documents contained in the Design Review packet and reconvene the Design Review Board at a later date.

    - The approved Initial Design Review packet which may contain the Design Control Statement, Requirements Document, Initial Risk Analysis and/or the Design and Development Plan as well as the approved Design Review information, will be placed in the Design History File.

NOTE:  In some instances, a formal Design Review may not be required.  In these instances, the Design Review packet could be reviewed by individual Design Review Board members without a formal meeting.  Additionally, the Initial and Final Design Review are outlined in this document.  It is likely that some changes or projects will require mid-phase Design Reviews.  These mid-phase Design Reviews are to be held and structured in a similar manner to the Design Reviews outlined in this document.

  - Once the Design Inputs have been approved, the development phase can be initiated.

    - Development results in the generation of production specifications as well as descriptive materials which define and characterize the design.

    - The Project Manager creates the design specifications as part of documenting design outputs and should include or reference all the relevant information required to manufacture, operate, maintain and market the device in accordance with its intended use and may refer to other device documentation including user manuals and assembly instructions.

    - The Project Manager collates the design outputs into a Design Input/output Matrix Form.  The design development process continues until all required inputs are met.  The Project Manager is responsible to update the matrix as the project progresses and the verification and validation activities demonstrate that the outputs satisfy the inputs.

    - Design reviews should be conducted where appropriate to support the design control activities.  Refer to Section E of FDA’s Design Control Guidance for Medical Device Manufacturers.

  - The next step of this phase is Design Verification.  Design Verification may include the testing, experimentation, analysis, calculations and/or evaluations as a confirmation that design outputs meet the design inputs.  Examples of Design Verification may include:

    - Evaluation of performance, stability, safety, reliability of the product or process under expected operational and storage conditions.

    - Inspections to verify that all design features are as intended and that design changes are identified and accomplished.

    - Biocompatibility testing of materials as applicable.

    - Bioburden/safety testing as applicable.

    - Measurements and/or dimensional analysis as applicable.

    - Analytical testing and analysis as applicable.

    - Equivalency testing to a design having a history of successful use as may be applicable.

    - Identify critical process control parameters and specifications as applicable.

    - Identification of material specifications and Qualification of raw materials as applicable.

    - Package integrity testing as applicable.

    - Literature Search/Scientific literature.

  - Design Validation and Validation activities are to be applied at various stages of the development as appropriate.

    - Design Validation is a cumulative summation of all efforts to assure that the product design will conform to the user/customer needs and intended use.  Design Validation is to target the use of the product in the actual or simulated environment, where possible.  Examples of Design Validation may include:

- Clinical trials/Clinical evaluation

- End user handling studies

- Actual or Simulated use environment tests

- Compilation of relevant scientific literature or provision of historical evidence

    - Process Validation activities are a cumulative summation of all efforts to assure the processes will conform to pre-determined requirements in the actual operating environment.  These Validation activities are intended to demonstrate that equipment, facilities, manufacturing processes and test methods perform as required.  Process Validation activities include:

- Installation Qualifications (IQ)

- Operational Qualifications (OQ)

- Performance Qualifications (PQ) such as:

- Aseptic processes

- Sterilization dose audits

- Software validations

- Process validations

    - Conduct required supplier qualification per appropriate SOP.

  - A Final Risk Analysis of the product or process is to be conducted.  Using the Initial Risk Analysis, an assessment is to be performed to determine if the risks in the Initial Risk Analysis were mitigated through Design Verification and Validation.  The failure modes, effects of a failure, cause of the failure and actual process controls are to be reassessed. The actual severity, occurrence and detection of failure are to be evaluated.  Based on this assessment, risk of failure can be assigned to product design or process aspects.  It is intended that the design verification and validation efforts significantly reduce the potential risk of product or process failures.

  - An Input/output Document may be prepared.  This document lists all the inputs and the outputs established throughout the design and development process for the new product or change to product or process.  References to verification and validation activities such as studies, experiments, analysis, tests, trials, validations and evaluations are appropriate.

  - Once the Design Outputs have been completed, a Final Design Review is to be convened.  This Final Design Review may include a review of the Design Verification activities, Design Validation and Validation activities, the Final Risk Analysis and the Input/output documents.

    - The Final Design Review is to include the following as appropriate:

- An agenda

- An objective

- Presentation of documents, results or other information pertinent to the Design Review

- Conclusions and/or Recommendations

- Outstanding items and assignments as appropriate

- Approvals

    - A Design Review packet should be prepared in advance and distributed to the Design Review Board.  The Board should review the Design Review packet and be prepared with comments, feedback, questions and any concerns.

    - Minutes and/or actions and assignments coming from the Final Design Review are to be documented.  Outstanding issues, if any, are to be addressed.  If outstanding actions from the Design Review Board cannot be closed out and addressed before the Design Review Packet is approved, these actions will be tracked as preventative actions according to CAPA system requirements.

    - If there are no issues or all issues can be resolved at the meeting, the Design Review Board will sign and date signifying ‘acceptance’ of the Final Design Review.

    - If all issues cannot be resolved at the meeting, actions will be identified and assignments made to resolve the issues.  Minor issues can be addressed and the Design Review packet can be circulated at a later time for approval.  If major issues are identified, it may be necessary to revise the documents contained in the Design Review packet and reconvene the Design Review Board at a later date.

    - The approved Final Design Review packet which may contain the Design Verification, Design Validation, Final Risk Analysis and/or the Input/output Document as well as the approved Design Review information, will be placed in the Design History File.

    - The appropriate output documents are to be submitted to Change Control for review and approval.  These document may include:

- Specifications

- Batch Records

- Testing Records

- Quality Control Records

- Drawings

- Labeling

- Inspection records

- Preventive Maintenance/Calibration records

NOTE:  In some instances, a formal Design Review may not be required.  In these instances, the Design Review packet could be reviewed by individual Design Review Board members without a formal meeting.  Additionally, new products or changes to existing process will be approved prior to the Design Transfer into Manufacturing.

  - Design Transfer is the bridging phase between design/development and product manufacturing.  It is a transfer of knowledge in the form of training, documentation and know-how.  This phase ensures that the Design Outputs have met the Design Inputs and the Design Outputs are correctly translated into production documentation. Design Transfer may include:

    - Training of Manufacturing and Quality Control employees and others to who the transfer is occurring.

    - Transfer of documentation for the manufacturing and testing of the product.  This documentation may include the following where applicable:

    - Bill of Materials (BOM)

    - Component fabrication details including material, tolerances, dimensions and finishing details

    - Assembly details including tolerances for critical interfaces

    - Calibration procedures

    - Device specifications (including limits of accuracy of measuring functions)

    - Instructions for use

    - Incoming inspection specifications and procedures

    - Electronic data files

    - Final inspection check list or procedure

    - Packaging and Labeling

    - Safety certification documentation

    - ESD/EMI/EMC certification documentation

    - Risk analysis documentation

    - Packaging test report

    - Preparation of a Device Master Record including all the documents required to manufacture a product.  The Device Master Record is reviewed and approved prior to final design transfer.  Follow EPIC’s Device Master Record procedure.

    - Transfer of the Design History records to Quality Assurance.

    - Verification that design specifications released to production have been approved, verified, and/or validated before they are implemented as part of the manufacturing process.

  - Design History File

    - The Project Manager is responsible to review documentation prior to inclusion in the DHF.  The DHF must contain all appropriate documentation to ensure that the necessary design controls have been completed and to fully characterize the design project and device development during the life cycle of the design history.

    - The DHF may be maintained and controlled by Engineering until the final design approval (i.e., technology transfer to Manufacturing) at which time all DHF documentation is placed in Document Control and Quality Assurance becomes responsible for maintaining and controlling the DHF.

    - The Design History File is to be compiled throughout the development process including records and documents that contain or reference all the design control elements required for the development of a new product or a change to an existing product or process.  All the documents outlined in the Design Control Statement are to be included in the Design History file. Once compiled, the Design History File is to be transferred to Quality Assurance and/or Document Control.  Documents that may be transferred as part of the Design History File may include:

- Design Control Statement

- Requirements Document

- Initial Risk Analysis Document

- Design and Development Plan

- Initial Design Review

- Mid-phase Design Reviews

- Design Verifications

- Design Validations

- Copies of or references to studies, experiments, evaluations, trials, tests, analysis or other documents that are a part of the development  process.

- Final Risk Analysis Document

- Input/output Document

- Final Design Review

- Design Transfer Documents/Records

- Device Master Record

    - DHFs will be maintained according to EPIC’s Records Retention and Destruction procedure.

  - The Design Control Process ends at this point.

  - Post Market/Manufacturing Surveillance:  Over time, changes or improvements to the product or production processes may be required.  Feedback from users/customer, Manufacturing or other groups may be obtained.  Based on the feedback, design changes may be required.  The proposed change can be considered and should be evaluated within the Design Control program as applicable.  See the Design Control Flow Chart.

- REFERENCES

21CFR820.30	Design Control, Quality System Regulations for Medical Devices

ISO13485	Quality System – Medical Devices – particular requirements for the application of ISO 9001

Design Control Guidance for Medical Device Manufacturers (FDA)

EG-010	Software Development

EG-012	Risk Management

EGF-001	Design Approval Form

EGF-002	Design Input and Output Matrix Form

EGF-003	Design Review Minutes Form

EGF-008	Risk Analysis Documentation Form

EGF-009	Failure Modes and Effects Analysis (FMEA)

EGF-010	Design Change Summary Form

EGF-012	Design Development Plan Template

OPS-008	Supplier Qualification

OPS-010	Records Retention and Destruction

QA-003	Change Control

QA-018	Corrective and Preventive Action Plans

- ATTACHMENTS

Attachment A, Medical Device Development Process Flow

Attachment A

# User/Customer Requirements/Essentials/Desirables

### Requirements Document

-----------------------------------------------

### Initial Risk Analysis Document

-----------------------------------------------

### Design and Development Plan

Initial Design Review

Acceptable?

Design Verification

Inputs=Outputs

----------------------------------------------

Design Validation & Validation Activities

----------------------------------------------

## Final Risk Analysis

## Input/Output Document

Final Design Review

Acceptable?

## Design Transfer

# Design Inputs

# Design Input Approvals

# Design Outputs

# Design Output Approvals

# Design History File

# Final Design History File to QA

#### No

#### Yes

#### No

Yes

# Change Existing Product or Process (Class I Change)

# New Product

#### Yes

#### Yes

## Post Market/ Manufacturing Surveillance

# Design Control Statement

## * Note: Specifications, batch records, SOP’s, drawings and other documents are to be submitted to Change Control

DESIGN

CONTROL

# Device Master Record

# To Change Control*

# To Change Control*

# Design History File

# Device Master Record

DESIGN

CONTROL

# Design Control Statement

#### Yes

#### Yes

# New Product

# Change Existing Product or Process (Class I Change)

Yes

#### No

#### Yes

#### No

# Final Design History File to QA

# Design Output Approvals

# Design Outputs

# Design Input Approvals

# Design Inputs

## Design Transfer

## Input/Output Document

Design Verification

Inputs=Outputs

----------------------------------------------

Design Validation & Validation Activities

----------------------------------------------

## Final Risk Analysis

### Requirements Document

-----------------------------------------------

### Initial Risk Analysis Document

-----------------------------------------------

### Design Development Plan

# User/Customer Requirements/Essentials/Desirables

# Concept & Feasibility

## * Note: Specifications, batch records, SOP’s, drawings and other documents are to be submitted to Change Control

Final Design Review

Acceptable?

Initial Design Review Acceptable?

What Design Control Elements Apply?

Document Revision History

| Change Order Number: | Version Number: | Description of Change: | Date: | Updated by: |
| --- | --- | --- | --- | --- |
| CC-052709-01 | 000 | New Document | 05/27/09 | TJ |
| CC-100809-01 | 001 | Update section 4.0 to reflect accurate responsibilities. Update section 7.0 to new format. Correct grammar/typographical errors throughout. | 10/08/09 | AM |
| CC-120601-01 | 002 | Complete rewrite of the entire procedure to combine the design control process into one comprehensive SOP. The title of the document was changed from Design Control Policy to Design Control Procedure. | 6/12/12 | A. Miller |