- Purpose

To develop a method of ensuring that the design requirements are translated into design outputs.  Furthermore, that design input requirements and output can be qualitatively or quantitatively compared.

- Scope

This procedure is applicable to approved product design and development projects for all new products or to major design changes in released products.

- Definitions

| Design Output | The results of a design effort at each design phase and at the end of the total design effort. The completed design output is the basis for the device master record. The total finished design output includes the device, packaging, labeling and DMR. |
| --- | --- |

- Responsibilities

| User | - Complies with the policy and procedure for all new or revised documents. - Ensures the most current version of this document is used when referenced. |
| --- | --- |
| Departmental Management | - Ensures departmental personnel are properly trained before using this policy or procedure. - Assigning personnel to design teams. - Reviews and approves all controlled documents for departmental procedures. |
| Quality Assurance | - Complies and enforces the document control standards within this document. - Monitors the implementation and effectiveness of this document. - Conducts periodic audits to assure compliance. |
| Engineering/Manufacturing | - Reviews and approves all products, product specifications, manufacturing processes, inspection processes, patient and physical manuals. - Reviews designs and procedures to ensure that inspection criteria can be performed. |
| Marketing | - Reviews the market and necessity for products. |

- Policy

  - Design outputs are summarized in the Design Input and Output Matrix.

  - Design outputs include production specifications as well as descriptive materials which define and characterize the design, which can include but are not limited to:

    - Assembly drawings

    - Component and material specifications

    - Production and process specifications

    - Software machine code

    - Work instructions

    - Quality Assurance specifications and procedures

    - Installation and servicing procedures

    - Packaging and labeling specifications

    - Risk analysis results

    - Software source code

    - Verification activities results

  - Design output procedures contain or make reference to specific acceptance criteria that ensure that those design outputs essential for the safe and proper functioning of the device (e.g. operating, storage, handling, maintenance and disposal requirements) are identified.

  - The design output is documented, reviewed and approved prior to release including the data and signature of the individual approving the output.

- Procedure

  - Based on the Design and Development Plan, the Project Manager summarizes the design outputs on the Design Input and Output Matrix.  Any contractors involved in the design process should be included in the Design Outputs.

  - The Project Manager creates the design specifications as part of the documented design outputs and should include or reference all the relevant information required to manufacture, operate, maintain and market the device in accordance with its intended use and may refer to other device documentation including user manuals and assembly instructions.  The device specification must include a detailed description or reference to the complete list of essential items identified in the Design Input and Output Matrix.

  - The Project Manager updates the matrix as the project progresses and the verification and validation activities demonstrate that the outputs satisfy the inputs.

  - Upon final completion, the matrix is reviewed by Quality Assurance and filed in the DHF.

- Reference

  - Design Input & Output Matrix, EGF-002

- Attachments

N/A

Document Revision History

| Change Order Number: | Version Number: | Description of Change: | Date: | Updated by: |
| --- | --- | --- | --- | --- |
| cc-052709-04 | 000 | New Document | 05/2709 | TJ |