- Purpose

This policy is intended to ensure that the design requirements are translated into design outputs.  Furthermore, that design input requirements and output can be qualitatively or quantitatively compared.

- Scope

This procedure is applicable to approved product design and development projects for all new products or to major design changes in released products.

- Definitions

| Design Output | The results of a design effort at each design phase and at the end of the total design effort. The completed design output is the basis for the device master record. The total finished design output includes the device, packaging, labeling and DMR. |
| --- | --- |

- Responsibilities

| User | - Complies with the policy and procedure for all new or revised product design. - Ensures the most current version of this document is used when referenced. |
| --- | --- |
| Departmental Management | - Ensures departmental personnel are properly trained before using this policy or procedure. - Assigning personnel to design teams. - Complies with the policy and procedure for all new or revised product design. - Ensures the most current version of this document is used when referenced. - . |
| Quality Assurance | - Complies and enforces the document control standards within this document. - Monitors the implementation and effectiveness of this document. - Conducts periodic audits to assure compliance. |
| Engineering/Manufacturing | - Reviews and approves all documentation related to device design development. - Participates in periodic design review to ensure that design controls are properly implemented. |
| Marketing | - Reviews the market and necessity for products. |

- Policy

  - Design outputs are summarized in the Design Input and Output Matrix.

  - Design outputs include production specifications as well as descriptive materials which define and characterize the design, which can include but are not limited to:

    - Assembly drawings

    - Component and material specifications

    - Production and process specifications

    - Software machine code

    - Work instructions

    - Quality Assurance specifications and procedures

    - Installation and servicing procedures

    - Packaging and labeling specifications

    - Risk analysis results

    - Software source code

    - Verification activities results

  - Design output procedures contain or make reference to specific acceptance criteria that ensure that those design outputs essential for the safe and proper functioning of the device (e.g. operating, storage, handling, maintenance and disposal requirements) are identified.

  - The design output is documented, reviewed and approved prior to release including the data and signature of the individual approving the output.

- Procedure

  - Based on the completed according to the Design and Development Plan, the Project Manager summarizes the design outputs on the Design Input and Output Matrix.  Any contractors involved in the design process should be included in the documentation of the Design Outputs.

  - Additionally, the final Design Input and Output Matrix will include the Design Inputs as established at the beginning of the design project.  Design Inputs may be modified with justification during the design development process.

  - The Project Manager creates the design specifications as part of documenting design outputs and should include or reference all the relevant information required to manufacture, operate, maintain and market the device in accordance with its intended use and may refer to other device documentation including user manuals and assembly instructions.  The device specification must include a detailed description or reference to the complete list of essential items identified in the Design Input and Output Matrix.

  - The Project Manager updates the matrix as the project progresses and the verification and validation activities demonstrate that the outputs satisfy the inputs.

  - Once completed (or at appropriate points throughout the design development), the Design Input and Output Matrix is reviewed by the design committee.  The review is documented on the Design Approval Form.  The form should document clearly approval or rejection of the Design Input and Output Matrix along with any required action items identified as a result of the review.

  - If contract suppliers are used during the any part of the design process, their input should be included into the Design Input and Output Matrix.

  - The design development process continues until all required inputs are met.  Once the final Design Input and Output Matrix has been approved, the manufacture process is then developed and a technology transfer from Engineering to Manufacturing occurs.  Design Verification and Validation must be completed prior to device distribution and a final Design Review completed and documented reviewing the entire Design History File (DHF) for compliance to quality system requirements.

  - All documentation will be maintained in the Design History File (DHF).  The DHF is maintained and controlled by Engineering until the final design approval (i.e., technology transfer to Manufacturing) at which time all DHF documentation is placed in Document Control and Quality Assurance becomes responsible for maintaining and controlling the DHF.

- Reference

EG-005, Design Review

EG-006, Design Verification and Validation

EGF-002, Design Input & Output Matrix

EGF-001, Design Approval Form

- Attachments

N/A

Document Revision History

| Change Order Number: | Version Number: | Description of Change: | Date: | Updated by: |
| --- | --- | --- | --- | --- |
| cc-052709-04 | 000 | New Document | 05/2709 | TJ |
| CC-100809-01 | 001 | Update section 4.0 to reflect current responsibilities. Add section 6.2 to provide clarification as to the inclusion of the previously developed Design Inputs in the final matrix. Add section 6.5 and 6.7 to provide committee review of the Design Input and Output Matrix and instructions on record storage. Update section 7.0 to reference all documents and revise format. | 10/08/09 | AM |