- Purpose

Verification and Validation activities are utilized to ensure that design input requirements are objectively embodied in the design outputs.  This procedure is intended to provide clarification for the control of documentation of test and device verification and validation activities.

- Scope

This procedure is applicable to approved product design and development projects for all new products or to major design changes in released products.

- Definitions

| Validation | Confirmation by examination and provision of objective evidence that the particular requirements for a specific intended use can be consistently fulfilled. |
| --- | --- |
| Process Validation | Establishing by objective evidence that a process consistently produces a result or product meeting its predetermined specifications. |
| Design Validation | Establishing by objective evidence that device specifications conform to user needs and intended uses. |
| Verification | Confirmation by examination and provision of objective evidence that specified requirements have been fulfilled. |
| Specification | Any requirement with which a product, process, service or other activity must conform. |
| Protocol | Document describing all activities to be completed in order to establish verification or validation of the design of a test and/or device. The protocol must include a matrix of responsibilities, the actual steps taken to establish verification and/or validation, and pass/fail criteria. |
| Final Report | The final report provides documentation of all activities and/or tests performed in the process of verifying and/or validating a test or device. This document must include a record of all personnel who performed the testing, the raw data collected as a part of testing, a summary of the results and a conclusion describing whether the verification/validation has been established. |

- Responsibilities

| User | - Complies with the policy and procedure for all new or revised design projects. - Ensures the most current version of this document is used when referenced. |
| --- | --- |
| Departmental Management | - Ensures departmental personnel are properly trained before using this policy or procedure. - Assigning personnel to design teams. - Complies with the policy and procedure for all new or revised product design. - Ensures the most current version of this document is used when referenced. |
| Quality Assurance | - Complies and enforces the document control standards within this document. - Monitors the implementation and effectiveness of this document. - Conducts periodic audits to assure compliance. |
| Engineering/Manufacturing | - Reviews and approves all documentation related to design verification and validation activities including protocols and final reports. - Responsible for planning, executing and summarizing design verification and validation activities. |

- Policy

  - Verification and validation activities must be planned, executed and documented and be representative of the complexity and risks associated with the product.

  - Software validation and verification activities are controlled through a separate procedure.

  - New processes or changes/deviations to existing EPIC processes are verified and/or validated prior to release of the products.

  - Validation is required to be performed and documented for process or design changes except when verification testing is acceptable.  Verification testing may be acceptable in such instances as minor component changes or for total label verifications.  The decision to perform validation versus verification testing is made on a case-by-case basis and is determined by reviewing the specific changes and its effect on form, fit, function, safety and effectiveness.

  - Final validation activities must be completed on devices representative of production units and their equivalency demonstrated through description photographs or other objective methods.

  - Initial production units fabricated by a manufacturing source must be validated to the approved specification for the device.

  - Validation is required for a design change that affects the intended use of the device.

  - All decisions regarding validation versus verification will be documented, reviewed and approved.

  - Approved validation and verification documents are maintained and filed in the design history file for the product being validated.

- Procedure

  - The Project Manager, in conjunction with the Design Review Committee, lists and approves the verification and validation activities in the Design and Development Plan and the Input and Output Matrix or a separate approved verification and validation plan created by the project manager.

  - Verification and validation activities may include the following:

    - Calculations

    - Risk controls identified in risk analysis

    - Testing to known standards or specifications or procedures

    - Comparative analysis

    - Testing

    - Interviews

    - Simulated testing

    - Extreme or destructive testing

  - Verification/validation activities are initiated once the design development team has established the Design Inputs are satisfied by the Design Outputs.  The first test of verification/validation activities is protocol development.  The protocol should be written according to the instructions provided in the Quality system Policy and Procedure Writing Guide.  The protocol will include the following:

    - A description of the device or component evaluated.

    - The test or challenge environment.

    - The test or challenge method including sampling plan with sample sizes and selection criteria.

    - A definition of the functional groups/individuals that is responsible for completing the verification/validation activities.

    - A complete description of required data collection including (where appropriate) records for data collection.

    - Responsibilities matrix to outline who is responsible for each task required in the validation.

    - A list of all required test equipment, including tolerances where applicable.

    - Acceptance criteria

  - The protocol is then executed.  The project manager is responsible to ensure that all aspects of the protocol are met or a deviation from the protocol obtained.  Product assembly is documented using the Device Master Record (DMR) or design documentation that fully characterizes the assembly.  All inspection and test results must be signed and dated by the individual recording the results.

  - Any deviations from the written protocol must be addressed and handled through the Deviations procedure.

  - A final report with a detailed table of contents outlining the verification/validation activities and summarizing all results is developed.

  - The final report should be written according to the instructions provided in the Quality system Policy and Procedure Writing Guide and contain the following :

    - A description of the device or component evaluated.

    - All required data collected including (where appropriate) data records, test equipment calibration records, non-conformance reports, approved deviations from protocol, and any other quality system records generated as a part of the validation activities.

    - Data analysis including a discussion of the data collected and a comparison to the acceptance criteria established in the protocol.

    - Conclusions including a clear statement of whether the validation/verification activities performed met the expected acceptance criteria

    - Follow up corrective actions, as applicable.

  - The final report document is routed for approval Protocol/Final Report Approval Process

    - The protocol and/or final report are routed for approval prior to execution/acceptance.

    - A Protocol/Final Report Approval Form is used as a coversheet to document approval.  Approval is required to be obtained for any deviation from the written protocol prior to execution.

    - Any unintended deviations from the written protocol are to be identified, discussed and reviewed as a part of the final report.  Unapproved deviations from protocol may result in rejection of the validation/verification activities at the discretion of the approval group.

  - Once completed, all documentation is stored in the Design History File (DHF) and/or (DMR).

  - The Project Manager monitors the verification and validation activities to ensure that testing and results are performed as described in the protocol.

  - Design Verification and/or Validation activities should be presented as a part of the final design review conducted prior to official release of the device manufacturing process.

  - The DHF is maintained and controlled by Engineering until the final design approval (i.e., technology transfer to Manufacturing) at which time all DHF documentation is placed in Document Control and Quality Assurance becomes responsible for maintaining and controlling the DHF.  The DMR is controlled through the Document Control system which is maintained by Quality Assurance.

- Reference

EG-002, Design and Development Planning

EG-004, Design Outputs

EG-016, Software Development Verification Testing Procedure

EGF-002, Design Input & Output Matrix

QA-001, Quality System Policy and Procedure Writing Guide

QA-004, Deviations

QA-025, Device Master Records

- Attachments

N/A

Document Revision History

| Change Order Number: | Version Number: | Description of Change: | Date: | Updated by: |
| --- | --- | --- | --- | --- |
| CC-052709-06 | 000 | New Document | 05/27/09 | TJ |
| CC-100809-01 | 001 | Update section 4.0 to reflect accurate responsibilities. Major update to section 6 to provide clarification of the verification/validation process and requirements. Update section 7.0 to revise format and reference correct policy/procedure/record. | 10/08/09 | AM |
| CC-120303-01 | 002 | Update procedure to outline responsibilities and requirements for managing and executing Verification and validation activities. Remove requirements to route verification and validation activities through EPIC’s Documentation Control System. | 3/5/12 | A. Miller |