- Purpose

This procedure provides instruction for the development and completion of software validation.

- Scope

Software validation is required to validate new or significantly modified software that is intended as a part of a finished medical device. This procedure is not intended to govern minor software “bug” fixes made to bring the software into compliance with documented expected functions.  This procedure is intended to be followed by all EPIC personnel who are involved with software development.

- Definitions

| Design Validation | Establishing by objective evidence that device specifications conform to user needs and intended uses. |
| --- | --- |
| Process Validation | Establishing by objective evidence that a process consistently produces a result or product meeting its predetermined specifications. |
| Requirements | The needs or expectations for a system and/or its software. There can be many different kinds of requirements (e.g., design, functional, implementation, interface, performance or physical requirements). |
| Specification | Any requirement with which a product, process, service or other activity must conform. |
| Verification | Confirmation by examination and provision of objective evidence that specified requirements have been fulfilled. |
| Validation | Confirmation by examination and provision of objective evidence that the particular requirements for a specific intended use can be consistently fulfilled. |

- Responsibilities

| User | - Complies with the policy and procedure for all new or revised design projects. - Ensures the most current version of this document is used when referenced. - Acts as project coordinator to ensure proper protocol development, accurate and timely protocol execution, and development of the final report. |
| --- | --- |
| Departmental Management | - Ensures departmental personnel are properly trained before using this policy or procedure. - Assigning personnel as project manager and/or to validation team(s). - Complies with the policy and procedure for all new or revised design projects. - Ensures the most current version of this document is used when referenced. - Reviews and approves protocols prior to validation and summary report of the validation activities to ensure that activities reflect actual activities performed, at a minimum. |
| Quality Assurance | - Complies and enforces the document control standards within this document. - Monitors the implementation and effectiveness of this document. - Conducts periodic audits to assure compliance. - Reviews and approves protocols prior to validation and summary report of the validation activities to ensure that all quality system requirements are met, at a minimum. |
| Engineering/Manufacturing | - Provides technical and specific guideance for protocol development and validation process execution. - Reviews and approves protocols prior to validation and summary report of the validation activities to ensure that summary report meets the technical requirements for software validation and that conclusions drawn are valid, at a minimum. |

- Policy

When/if computers or automated data processing systems are used as part of a finished device or a part of EPIC’s quality system, the software used is validated for its intended use.  New processes or changes/deviations to existing software are verified and/or validated prior to release of the products.  Validation requirements are reviewed and addressed for each product to ensure adherence to all proposed uses and functions according to applicable specifications and requirements.  Validation is required to be performed and documented for process or design changes except when verification testing is acceptable.  Verification testing may be acceptable in such instances as minor code changes or where changes can be 100% verified.  The decision to perform validation versus verification testing is made on a case-by-case basis and is determined by reviewing the specific changes and its effect on form, fit, function, safety and effectiveness.  All decisions regarding validation versus verification will be documented.  Approved validation and verification documents are maintained and filed in Quality Assurance.

- Procedure

  - Quality Assurance users will be notified by the Gemini (Bug tracking and Enhancement Management Software) automatically when a bug or enhancement has been flagged as completed. This will allow the QA team to align testing resources and develop a test plan based on the type of change that was made to the software.

  - Quality Assurance will coordinate with development to get the test machines loaded with the new version of the software. This may not happen every time a new bug fix or enhancement is made.

  - Quality Assurance will mark the completed item as “QA Testing” in the Gemini system before beginning the testing process.

Quality Assurance will compare the behavior of the change to the documentation of the change contained in Gemini. If all requirements are met, the tester will then mark the change as ‘Passed QA’  If the requirements are not met, a comment should be added to the item explaining the variance and the status is changed to ‘Reopened’. The development will then revisit the issue, correct the problem and then mark the item as needing QA again.  QA personnel are required to print out the any Gemini submission, sign and date the original copy and forward to Quality Assurance for filing within the Design History File.

  - The release will remain pending until all of the items defined for the release have a status of ‘Passed QA’. At this point, a regression test will be done.  Regression testing will be performed according to written protocols released for use.  If a protocol has not been established, one will be developed, reviewed and approved by the development team prior to execution.  Information Systems staff will work with Quality Assurance staff to develop the protocols.  If regression testing finds any issues, a new Gemini issue should be opened for the problem which will prompt the development cycle to begin again.

  - If the statuses for the items remains at ‘Passed QA’ and no new items have been added from the regression testing, the version will be deemed ready for release. At this point, Development will set a marker in the version control system denoting the master copy of that version of software.

- Reference

EG-010, Software Development

- Attachments

N/A

Document Revision History

| Change Order Number: | Version Number: | Description of Change: | Date: | Updated by: |
| --- | --- | --- | --- | --- |
| CC-011810-01 | 000 | New Document | 1/26/2010 | AFM |