- Purpose

This procedure provides instruction for the development and completion of software validation.

- Scope

Software validation is required to validate new or significantly modified software that is intended as a part of a finished medical device. This procedure is not intended to govern minor software “bug” fixes made to bring the software into compliance with documented expected functions.  This procedure is intended to be followed by all EPIC personnel who are involved with software development.

- Definitions

| Design Validation | Establishing by objective evidence that device specifications conform to user needs and intended uses. |
| --- | --- |
| Process Validation | Establishing by objective evidence that a process consistently produces a result or product meeting its predetermined specifications. |
| Requirements | The needs or expectations for a system and/or its software. There can be many different kinds of requirements (e.g., design, functional, implementation, interface, performance or physical requirements). |
| Specification | Any requirement with which a product, process, service or other activity must conform. |
| Verification | Confirmation by examination and provision of objective evidence that specified requirements have been fulfilled. |
| Validation | Confirmation by examination and provision of objective evidence that the particular requirements for a specific intended use can be consistently fulfilled. |

- Responsibilities

| User | - Complies with the policy and procedure for all new or revised design projects. - Ensures the most current version of this document is used when referenced. - Acts as project coordinator to ensure proper protocol development, accurate and timely protocol execution, and development of the final report. |
| --- | --- |
| Departmental Management | - Ensures departmental personnel are properly trained before using this policy or procedure. - Assigning personnel as project manager and/or to validation team(s). - Complies with the policy and procedure for all new or revised design projects. - Ensures the most current version of this document is used when referenced. - Reviews and approves protocols prior to validation and summary report of the validation activities to ensure that activities reflect actual activities performed, at a minimum. |
| Quality Assurance | - Complies and enforces the document control standards within this document. - Monitors the implementation and effectiveness of this document. - Conducts periodic audits to assure compliance. - Reviews and approves protocols prior to validation and summary report of the validation activities to ensure that all quality system requirements are met, at a minimum. |
| Engineering/Manufacturing | - Provides technical and specific guideance for protocol development and validation process execution. - Reviews and approves protocols prior to validation and summary report of the validation activities to ensure that summary report meets the technical requirements for software validation and that conclusions drawn are valid, at a minimum. |

- Policy

When/if computers or automated data processing systems are used as part of a finished device or a part of EPIC’s quality system, the software used is validated for its intended use.  New processes or changes/deviations to existing software are verified and/or validated prior to release of the products.  Validation requirements are reviewed and addressed for each product to ensure adherence to all proposed uses and functions according to applicable specifications and requirements.  Validation is required to be performed and documented for process or design changes except when verification testing is acceptable.  Verification testing may be acceptable in such instances as minor code changes or where changes can be 100% verified.  The decision to perform validation versus verification testing is made on a case-by-case basis and is determined by reviewing the specific changes and its effect on form, fit, function, safety and effectiveness.  All decisions regarding validation versus verification will be documented.  Approved validation and verification documents are maintained and filed in Quality Assurance (QA).

- Procedure

  - IT completes the development package and forwards all required documentation after completing verification testing as outlined in the Software Development Verification Testing Procedure.  This package serves as notification to QA that software is ready for the software validation process.

  - QA will review the provided documentation for compliance to all requirements and develop a test plan based on the type of change that was made to the software.

  - Quality Assurance will coordinate with development to get the test machines loaded with the new version of the software.

  - The software version is controlled by IT during testing to ensure proper version tracking as corrections are made during the validation test process.  Each iteration will receive an incremented value in the version to denote changes that may be summarized at the end of the validation test process.  IT will maintain a brief list of these changes as a part of the versioning process and will forward a final complete list to QA at the end of validation testing for inclusion in the validation analysis.

  - Upon completion of the validation testing, a summary of all test plans, expected outcomes, and test results including the summary of any version changes executed during development will be documented as a validation analysis to be included in the final design review.

- Reference

EG-010, Software Development

EG-016, Software Development Verification Testing Procedure

- Attachments

N/A

Document Revision History

| Change Order Number: | Version Number: | Description of Change: | Date: | Updated by: |
| --- | --- | --- | --- | --- |
| CC-011810-01 | 000 | New Document | 1/26/2010 | AFM |
| CC-120202-01 | 001 | Rewrite procedure to remove reference to tracking in the Gemini system and reflect current design development activities. | 2/3/12 | A. Mason |