- Purpose

This procedure provides instruction for the development and completion of software validation.

- Scope

Software validation is required to validate new or significantly modified software that is intended as a part of a finished medical device. This procedure is intended to be followed by all EPIC personnel who are involved with software development.

- Definitions

| Design Validation | Establishing by objective evidence that device specifications conform to user needs and intended uses. |
| --- | --- |
| Process Validation | Establishing by objective evidence that a process consistently produces a result or product meeting its predetermined specifications. |
| Requirements | The needs or expectations for a system and/or its software. There can be many different kinds of requirements (e.g., design, functional, implementation, interface, performance or physical requirements). |
| Specification | Any requirement with which a product, process, service or other activity must conform. |
| Verification | Confirmation by examination and provision of objective evidence that specified requirements have been fulfilled. |
| Validation | Confirmation by examination and provision of objective evidence that the particular requirements for a specific intended use can be consistently fulfilled. |

- Responsibilities

| User | - Complies with the policy and procedure for all new or revised design projects. - Ensures the most current version of this document is used when referenced. - Acts as project coordinator to ensure proper protocol development, accurate and timely protocol execution, and development of the final report. |
| --- | --- |
| Departmental Management | - Ensures departmental personnel are properly trained before using this policy or procedure. - Assigning personnel as project manager and/or to validation team(s). - Complies with the policy and procedure for all new or revised design projects. - Ensures the most current version of this document is used when referenced. - Reviews and approves protocols prior to validation and summary report of the validation activities to ensure that activities reflect actual activities performed, at a minimum. |
| Quality Assurance | - Complies and enforces the document control standards within this document. - Monitors the implementation and effectiveness of this document. - Conducts periodic audits to assure compliance. - Reviews and approves protocols prior to validation and summary report of the validation activities to ensure that all quality system requirements are met, at a minimum. |
| Engineering/Manufacturing | - Provides technical and specific guidance for protocol development and validation process execution. - Reviews and approves protocols prior to validation and summary report of the validation activities to ensure that summary report meets the technical requirements for software validation and that conclusions drawn are valid, at a minimum. |

- Policy

When/if computers or automated data processing systems are used as part of a finished device or a part of EPIC’s quality system, the software used is validated for its intended use.  New processes or changes/deviations to existing software are verified and/or validated prior to release of the products.  Validation requirements are reviewed and addressed for each product to ensure adherence to all proposed uses and functions according to applicable specifications and requirements.  Validation is required to be performed and documented for process or design changes except when verification testing is acceptable.  Verification testing may be acceptable in such instances as minor code changes or where changes can be 100% verified.  The decision to perform validation versus verification testing is made on a case-by-case basis and is determined by reviewing the specific changes and its effect on form, fit, function, safety and effectiveness.  All decisions regarding validation versus verification will be documented.  Approved validation and verification documents are maintained and filed in the appropriate design history file (DHF) maintained by Quality Assurance (QA).

- Procedure

  - IT completes software construction and unit testing as outlined in the Software Development Verification Testing Procedure.  IT will then work with QA to generate a software verification and validation plan.

  - Once developed, the software verification and validation plan will be reviewed and approved.  Approval will be obtained from all functional groups who have participated as design reviews of the requirements, specifications, and risk analyses.  The approval group will contain the following:

    - Executive Management

    - IT

    - QA

    - Clinical/Marketing (if software is being used in product intended for sale)

    - Manufacturing (if software is being used in product assembly/manufacture)

    - Engineering (or appropriate technical expert)

  - Quality Assurance will coordinate with development to get the test machines loaded with the new version of the software.  IT will document the test equipment set-up.  This documentation will be included in the final summary report.

  - The software version is controlled by IT during testing to ensure proper version tracking as corrections are made during the validation test process.  Each iteration will receive an incremented value in the version to denote changes that may be summarized at the end of the validation test process.  IT will maintain a brief list of these changes as a part of the versioning process and will forward a final Revision History Table to QA at the end of validation testing for inclusion in the final validation analysis/report.

  - QA will generate appropriate verification and validation protocols based on the approved software verification and validation plan.  These protocols will be developed and approved using EPIC Engineering Studies procedure.  Approval will include the same functional groups as included in the approval of the software verification and validation plan unless otherwise justified in writing to the design history file.

  - Protocol Execution

    - QA will coordinate execution of the software verification and validation test protocols.  All records will be retained and included in the final summary report.

    - QA and IT will coordinate the investigation and resolution of any non-conformances and/or deviations that surface during validation testing.  If code changes are required to resolve non-conformances, IT will issue a new revision and modify the code which will then be deployed for use in retesting.

  - Final Software Validation Test Report

    - Upon completion of the validation testing, a summary of all test plans, expected outcomes, and test results including the summary of any version changes executed during development will be documented in a summary report using EPIC’s Engineering Studies procedure.

    - A validation analysis may to be included in the final report to establish documentation to summarize multiple validation test results.

    - The final report approval will include the same functional groups included in the test protocol approval.

  - Upon approval of the final report of all verification and validation tests, QA will forward the final reports to the project manager for inclusion in the final  design review.

- Reference

EG-010, Software Development

EG-016, Software Development Verification Testing Procedure

EG-020, Engineering Studies

- Attachments

N/A

Document Revision History

| Change Order Number: | Version Number: | Description of Change: | Date: | Updated by: |
| --- | --- | --- | --- | --- |
| CC-011810-01 | 000 | New Document | 1/26/2010 | AFM |
| CC-120202-01 | 001 | Rewrite procedure to remove reference to tracking in the Gemini system and reflect current design development activities. | 2/3/12 | A. Mason |
| CC-120514-02 | 002 | Rewrite procedure to reflect current practices and clearly indicate order of operations. Add reference to new Engineering Studies procedure. | 5/14/12 | A. Miller |