- Purpose

To develop a method of identifying, evaluating and prioritizing risks related to product development, product quality, product safety and product changes.

- Scope

This procedure is applicable to approved product design and development projects for all new products or to major design changes in released products.

- Definitions

| Failure Mode and Effects Analysis (FMEA) | A procedure by which each potential failure mode in a system is analyzed to determine the results or effects on a system and to classify failure modes according to severity. FMEA is a bottom-up system safety analysis method. |
| --- | --- |
| Fault Tree Analysis (FTA) | An analytical technique whereby an undesired system state is specified and the system is then analyzed to find all credible ways in which the undesired event can occur. FTA is a top-down system safety analysis method. |
| Risk Priority Number (RPN) | An expression of the possibility of a mishap in terms of probability and severity. |
| Safety | Freedom from conditions that can cause death, injury, or occupational illness. May also include damage to or loss of equipment or property. |
| Safety Related Software | Software functions whose errors can result in a potential safety mishap or loss of predictability or control. |
| Safety Engineering | An engineering discipline requiring specialized professional knowledge and skills in applying scientific and engineering principles, criteria and techniques to identify and eliminate or reduce safety skills. |
| Safety Risk Analysis | A set of procedures to evaluate safety risks and define safety design criteria. The safety risk analysis should be started during the earliest life-cycle phases of development. |
| Severity | A severity rating is defined to provide a measure of the worst credible mishap that could result should a safety mishap occur. The severity rating takes into consideration the defined safety requirements. Severity is based on a qualitative measure of the potential consequence of a safety hazard and is defined to be one of the categories as shown in Table 1. |
| Occurrence | Occurrence is defined as the likelihood of occurrence over the life of the system taking into consideration the defined safety requirements. For example, in determining occurrence, the total user population, hours of operation and system life are to be taken into account. |
| Detection | Detection is defined as the assessment of the existing/proposed controls to identify any potential failure mode prior to occurring or prior to release of the product for manufacture and/or distribution. |

- Responsibilities

| User | - Complies with the policy and procedure for all new or revised documents. - Ensures the most current version of this document is used when referenced. |
| --- | --- |
| Departmental Management | - Ensures departmental personnel are properly trained before using this policy or procedure. - Assigning personnel to design teams. |
| Quality Assurance | - Complies and enforces the document control standards within this document. - Monitors the implementation and effectiveness of this document. - Conducts periodic audits to assure compliance. |
| Engineering/Operations | - Reviews and approves all products, product specifications, manufacturing processes, inspection processes, patient and physical manuals. - Reviews designs and procedures to ensure that inspection criteria can be performed. |
| Marketing | - Reviews the market and necessity for products. |

- Policy

  - The Risk Analysis is intended to be a living document throughout the design cycle and into the manufacture of finished devices.  The document is intended to be used as appropriate in ongoing quality system reviews including nonconforming product disposition, corrective and preventive action decisions, and change control approval.

  - A design team made up of key participants (e.g., engineering, clinical staff, marketing, users, regulatory personnel, operations, and/or quality assurance) will be responsible for development of the risk analysis.  The design team will be responsible to conduct a risk assessment at appropriate phases of the design project utilizing the product specifications, design inputs and design outputs as appropriate.

  - After release of the initial risk assessment, the design team and/or senior management will have the responsibility to ensure that significant changes to the product specifications, manufacturing process, and/or intended use of the finished product are assessed for potential risks and any appropriate modifications to the risk assessment are made through the change control process.

  - The corrective and preventive action system will be utilized to identify potential areas of modifications to the risk assessment based on the historical data of nonconformances, complaints from users, and areas of required corrective and preventive action.  This evaluation will be conducted as a minimum during the annual Management Review.

- Procedure

  - The design team will gather appropriate information regarding the potential hazards associated with using the system/device.  This information will be reviewed for use in the documentation of the risk assessment.  Key documents might include (as available):

- Product Specification

- Design Input and Output Matrix

- Instructions for Use

- Regulatory documentation (e.g., 510k, PMA, etc.)

- Process flowcharts

- Complaint files

- Nonconforming material reports

- Clinical studies, animal studies, and other scientific data

NOTE:  For an initial or preliminary risk analysis, there may be limited information available to review prior to the meeting.

  - Identification of Risks

The design team may select one individual to facilitate the documentation process to ensure that all identified potential areas of risk are assessed whether or not they are subsequently added to the final risk assessment documentation.

    - Identification of Qualitative and Quantitative Characteristics- Utilize the Risk Analysis Documentation Form to document the qualitative and quantitative characteristics of the system/device.

    - Identification of Possible Hazards-  Any characteristic listed under Step 2 of the Risk Analysis Documentation Form will be utilized to identify hazards to the user and/or patient that could occur under normal use conditions or due to a single fault.

NOTE:  The design team may use section 3 of IEC 60513 to assist in identifying possible hazards.

  - Estimation and Assessment of the Risks for Each Hazard

    - Meeting minutes and/or other documentation (e.g., e-mail review/comments) should be generated to support the content/decisions represented in the Failure Modes and Effects Analysis (FMEA) analysis.

    - The Failure Modes and Effects Analysis (FMEA) Form (or equivalent) will be utilized to document the estimation and assessment of risks for each hazard.  Fill out the required information to identify the project/device.

    - Review each possible hazard from the Risk Analysis Documentation Form and document each column of information as appropriate for the hazard being reviewed.

      - Each part number/device listed in column 1 will be individually numbered sequentially.  Complete each column identifying hazards and possible accidental events that can reasonably be expected to occur.  The assessment should consider all parts of the system, operational modes, maintenance modes, safety systems, etc.

      - Hazards may be categorized into Device Hardware, Device Software, and Use Related.  The following potential defects can contribute to risk and will be assessed for each category:

Device Hardware and Software

    - Design Defects

    - Manufacturing Defects

    - Warning Defects

Use Related (by healthcare professional and/or patient)

    - Avoidable Known Improper Conditions of Use

    - Unavoidable Known Improper Conditions of Use

    - Unknown consequences

      - Associated with long term use

      - Associated with concurrent use of other products

      - Associated with off-label (unapproved or unstudied) uses

      - Associated with use in an unstudied patient population

    - The Design team documents the risk assessment

      - Document the part name and number and the process/product function to be assessed.

      - Identify and document potential failure modes associated with the system/device and function.  If the process/product function has a low risk, document rationale why no analysis is necessary.

      - Document the potential effects of the failure and select a severity.  Severity is an assessment of the seriousness of the potential effect of failure on the local level, higher levels, and the end user/patient.  Changes in the severity rating should only occur when there is a design change or an application change.

All possible causes of a failure mode having the same consequence should be assigned the same severity ranking.  Assign a severity ranking that represents the degree of severity of the potential effect of failure for each reviewed cause from the following table:

Table 1

| Ranking | Degree of Severity |
| --- | --- |
| 1, 2 | The failure is a cosmetic defect only. The failure will not have any perceptible effect on the performance of the product. The user will probably not notice the failure. |
| 3, 4 | The user is only minimally affected. The user will probably notice only a minor deficiency in the product or system. |
| 5 | The failure causes dissatisfaction on the part of the one affected. The user will notice decreased product or system performance. The product is operable at reduced performance level. The use of the device may be prolonged and the patient interface with the device may be elongated. |
| 6 | The failure necessitates medical or surgical intervention by a health care professional. |
| 7 | The failure necessitates medical or surgical intervention by a health care professional to preclude permanent damage to body function or permanent damage to the body structure. |
| 8 | The failure results in permanent impairment of the body function or permanent damage to a body structure. |
| 9 | The failure results in a complication that is life threatening. |
| 10 | The failure results in a complication that would lead to death. |

      - Determine the potential cause of failure and assess the probability of occurrence.

Occurrence is the likelihood that a specific cause will result in a specific failure mode.  The following questions can be used in assess the occurrence.

- Are the current controls effective?

- Is the product/process/assembly/part a carryover from a previous or similar design?

- How significant are the changes from the previous design?

- What are the environmental changes?

- What does the warranty or complaint data confer?

- What are the customer’s expectations?

Assign a ranking that represents the likely occurrence rate of the reviewed cause using the following table.

Table 2

| Ranking | Occurrence Rating |
| --- | --- |
| 1 | Remote, failure unlikely |
| 2, 3 | Low, relatively few failures |
| 4, 5, 6 | Moderate, occasional |
| 7, 8 | High, repeated failures |
| 9, 10 | Extreme, almost inevitable failure |

      - Consider the process controls that exist or could exist to mitigate the severity or occurrence.  Document all know process controls.  Determine the effectivity of the process control to detect the failure and record the number in parenthesis next to the listed process control.  Record the highest detection number in the Detection column.

Detection is the assessment of the existing/proposed controls to identify any potential failure mode prior to occurring or prior to release the product for production or distribution.  Rating of the probability of detection is based on the effectiveness of the existing/proposed control system throughout the chain of design events and/or current manufacturing/QC controls.  To receive a better rating, either the existing/proposed control system must be improved or the design must be revised to improve the effectiveness of the current control.

Assign a ranking that represents the effectiveness of the detection of the controls for each reviewed cause and subsequent failure mode from the following table.

Table 3

| Ranking | Effectively of Detection |
| --- | --- |
| 1, 2 | Controls will almost certainly detect or enable the user to detect a potential cause and subsequent failure mode. Current control(s) will almost certainly detect the failure mode. Reliable detection controls are known in similar process. |
| 3, 4 | Very likely that controls will detect a potential cause and subsequent failure mode prior to release to inventory. Very high likelihood that current controls will detect failure mode and/or control will enable user to detect a potential cause and subsequent failure mode. |
| 5, 6 | Moderate to low likelihood that controls will detect or enable the user to detect a potential cause and subsequent failure or that current controls will detect the failure mode. |
| 7, 8 | Remote chance that controls will detect or enable the user to detect a potential cause remote likelihood that current controls will detect failure. |
| 9, 10 | Undetectable until failure occurs in the field or no control available to detect the failure mode. |

      - Compute the Risk Priority Number (RPN)-  The risk priority number is a weighted assessment of the causes of failure.  It is comprised of the occurrence, severity, and detection levels.  The RPN can be achieved by multiplying all three components together.  A comparison of the values calculated should assist in the prioritization of the work associate with minimizing failures, safety issues, or meeting customer requirements.

The RPN will be a number between 1 and 1000.  The guideline for the maximum acceptable RPN value is 216 (RPN = 6 x 6 x 6 = 216).  Regardless of the RPN value, severity or detection ratings above six (6) require additional mitigation to reduce their rating.  If no additional changes are recommended to decrease the severity and/or detection ratings in these cases, then a justification for the decision must be entered into the Rationale column on the form.

      - Mitigate any unacceptable RPN, severity and/or detection ranking by eliminating, correcting, minimizing, or accepting (through documented rationale) the defects or weaknesses by redesigning, reformulating, labeling, etc.  Any newly instituted controls should be referenced in the New Controls section.  Reference should be made to policies/procedure document numbers, etc. which are instituted to mitigate potential risk.  Mitigation of the hazards should continue until the RPN, Severity and/or Detection levels are acceptable or a rationale for the level is documented.

  - After all potential risks have been collected on the Risk Analysis Documentation Form and the FMEA form; the forms will be routed through the Documentation System Policy and managed under change control as independent documents for each product/device.

  - The risk analysis is intended to be utilized to evaluate future changes to product design and or manufacturing process.  Proposed device changes, nonconforming history trend analysis, corrective/preventive action trend analysis and/or complaint history records will be evaluated to determine if new potential risks are created/identified and/or if the suggested changes increase or decrease the potential seriousness of risk and severity of potential harm.  Any new potential risks (failure modes, causes and/or effects) will be evaluated and documented as required in this procedure.

- Reference

EGF-002, Design Input and Output Matrix

EGF-008, Risk Analysis Documentation Form

EGF-009, Failure Modes and Effects Analysis (FMEA) Form

IEC 60513:1994, Fundamental Aspects of Safety Standards for Medical Electrical Equipment

IEC 60601-1-4:2000, Medical Electrical Equipment- Part 1-4 General Requirements for Safety- Collateral Standard: Programmable Electrical Medical Systems.

EN 1441, Risk Analysis: Medical Devices

- Attachments

N/A

Document Revision History

| Change Order Number: | Version Number: | Description of Change: | Date: | Updated by: |
| --- | --- | --- | --- | --- |
| CC-091409-01 | 000 | New Document | 09-21-09 | AM |