- Purpose

This procedure provides policy and procedure for testing that must be done by the developers prior to the changes being submitted to quality for validation and verification.

- Scope

This procedure applies to the testing of all functionality added to the software.  This procedure is intended to be followed by all EPIC personnel who are involved with software design and development.

- Definitions

| Unit Testing | Testing of a new component or aspect of the software. This is verification that the process that was described in the requirements documents is performed by the new code that was created. |
| --- | --- |
| System/Integration Testing | Testing that is performed to verify that the new component or functionality that was created is properly accessible through the system and produces the expected results when run through the system. |
| Non-Breaking Change | A change to the software that is purely cosmetic, not affecting any of the business or calculation logic. |
| Requirements | The needs or expectations for a system and/or its software. There can be many different kinds of requirements (e.g., design, functional, implementation, interface, performance or physical requirements). |
| Specification | In software development, a specification is defined as a document that states the requirements |
| Verification | Confirmation by examination and provision of objective evidence that specified requirements have been fulfilled. |
| Validation | Confirmation by examination and provision of objective evidence that the particular requirements for a specific intended use can be consistently fulfilled. Also said as, confirmation by examination and provision of objective evidence that software specifications conform to use needs and intended uses, and that the particular requirements implemented through software can be consistently fulfilled. |

- Responsibilities

| User | - Complies with the policy and procedure for all new or revised projects. - Ensures the most current version of this document is used when referenced. |
| --- | --- |
| Departmental Management | - Ensures departmental personnel are properly trained before using this policy or procedure. - Assigning personnel as project manager and/or to validation team(s). - Complies with the policy and procedure for all new or revised design projects. - Ensures the most current version of this document is used when referenced. - Reviews and approves Software Verification Records. |
| Quality Assurance | - Complies and enforces the document control standards within this document. - Monitors the implementation and effectiveness of this document. - Conducts periodic audits to assure compliance. - Reviews and approves Software Verification Records. |

- Policy

EPIC is committed to performing unit testing (verification testing) of software during software construction.  This unit testing is intended to focus on early examination of the sub-program functionality and ensure that the functionality not visible at the system level is examined by testing.  Unit testing ensures that quality software units are furnished for integration into the finished software project.  EPIC is committed to documenting all test methods and data (results) to ensure that objective pass/fail decisions are made.  Software verification test results should be used to assess the preliminary risk analysis produced during earlier development phases and update the documented risks as necessary during the software construction process.

- Procedure

  - When a requirement or logical set of requirements have been implemented, it will be necessary for the developer to  perform unit and system/integration testing prior to the code being checked into the source code control system. This is done at the end of the implementation process; testing of iterative changes done during development may be conducted but are not subject to the documentation requirements outlined this procedure.

  - A Software Verification Record will be completed prior to the code being checked into the source code control system. This form requires the developer to verify that basic good practices (as outlined in the form) have been completed. In addition the form contains an open text area for the developer to write notes specific to the testing that was completed. One to many Software Verification Records can be used during the testing process, failed tests should be kept for record as well as passing ones.

    - The Requirement section of the record is used to record the reference number for the requirement being tested.  The reference number is located in the software requirements document.

    - The Expected Outcome section of the record can be a record of the specifications, but the minimum information required to be recorded must identify the expected software performance requirement.

For example:  If the software requirement is to store a data point in a numerical format to two decimal points, the developer could record “Datapoint X saves numerical data to two decimal points.  Developer entered a “1” with no decimal points into the X datapoint.”

    - The Outcome column of the record is used to record the actual observed outcome of the test identified in the Expected Outcome column.

For example:  Using the expected outcome specified in the section above, a passing verification test would confirm that the database stored a “1.00” for databpoint X.

    - The developer will initial and date in the spaces provided after completing each verification test.

    - Some software requirements may require that more than one verification test to ensure proper functionality.

For example: Using the example specified above.  An additional two verification tests would be required to be performed in this case.  The second verification test would require entering “1.11” in order to confirm that the correct data is stored correctly.  The Expected Outcome of this second test could be recorded as “Enter and save 1.11 into datapoint X.  Verify that 1.11 is stored for datapoint X in the database.” The third verification test would require entering “1.123” in order to confirm that the data is stored correctly.  The Expected Outcome of this second test could be recorded as “Enter and save 1.123 into datapoint X.  Verify that 1.12 is stored for datapoint X in the database.”   In some cases these verifications can be combined, but the expected outcome and actual outcome should be clearly delineated to ensure that the record establishes objective evidence that the software constructed meets the established requirements.

    - Once all verification tests are recorded, the developer will complete their name, department, sign, and date the form in the space provided.  The developer will forward the record to the designated IT staff for review and approval.

    - IT staff will review the completeness, accuracy and appropriateness of all verification testing documented on the Software Verification Record.  If all record keeping requirements are met, the IT designee will sign and date the record in the space provided.  The IT designee will then forward the Software Verification Records to a QA staff designee for review and filing.

    - The QA designee will review the Software Verification Record for compliance to this procedure.  If any issues are surfaced, the QA designee will work with the developer and the IT designee to complete the record before final approval and filing.  Once all issues are resolved, the QA designee will sign and date the form in the space provided and file the record in the project design history file.

  - Developers may cycle into various phases of verification tests during software construction.  When all testing has been completed with passing results (or at appropriate intervals), the code should be checked into the source code control system.

  - Refer to the specific DDP, project plan and the Software Development procedure for next steps in the development process.

- Reference

EG-010, Software Development

EGF-011, Software Verification Record

- Attachments

N/A

Document Revision History

| Change Order Number: | Version Number: | Description of Change: | Date: | Updated by: |
| --- | --- | --- | --- | --- |
| CC-120202-01 | 000 | New Document | 2/3/12 | A. Mason |
| CC-120328-01 | 001 | Update the policy section to match current verification testing expectations. Update section 6.2 to provide clarification and examples on the record keeping requirements. Remove requirements for next steps and refer user to the Software Development SOP. | 5/14/12 | A. Mason |