# Purpose

This document describes the Software Code Standards that EPIC Research & Diagnostics, Inc. will use in developing the embedded software for the ClearView device.  The standards and guidelines described in this document are expected to be adhered to for all software developed for the embedded portion of the ClearView Device Software.

# Scope

This document was written with the following goals in mind:

- Code should be robust, deterministic and error free.

- Code should be readily understandable, readily  unit-tested, and readily used.

- Code should be easy to peer-review and maintain.

- Code should be readily verifiable.

Each section of this document contains standards in the corresponding area, along with a further list of recommended guidelines. These standards are required to be followed in all current and future development. If a programmer must deviate from a standard, then the reason for doing so must be strong and clearly documented via code comments. Individual developers are encouraged to program in a style that is consistent with the rest of the development team, but differing code styles are acceptable so long as they adhere to these coding standards and the corresponding code review checklist.  Code should enforce the concept of team-standardization, which precludes code from being so unique as to be readily identified with being originated by a single author; this Standard provides the details to ensure such.

# Document Overview

The purpose of this document is to describe the EPIC Research & Diagnostics, Inc, standard for software coding activities.  The intent of this document is to provide a consistent and concise definition of this standard for use within the coding phase of all embedded software engineering projects at EPIC Research & Diagnostics, Inc..

# Coding Standards for c

## Standards

All code should be ANSI standard and should compile without warning under at least its principal compiler. Any warnings that cannot be eliminated should be commented in the code.  All code should be reviewed by its author, and an independent reviewer (any competent engineer aside from the originator)

# Naming Convention

One of the best means for creating software that is easily understandable is by making all names clear, unambiguous, and informative. It is imperative to avoid abstract names (in a global context) that are likely to be reused by other parts of the system

## Module Names

### Standards

File names should represent the content or role of the file.

  - Header file names shall have the extension “.h”.

  - C Implementation (source) file names shall have the extension “.c”

  - File names shall contain only alphanumeric characters and “_” (underscore). No other characters are permitted.

Mixed case names are permitted when only the first letter of an Identifiable word is in Upper case

Example:

MyFirstModuleName.c

MyFirstModuleName.h

My_First_Module_Name.c

My_First_Module_Name.h

## Function Names

### Standards

Function names will be unique and identify as far as possible the action performed or the information provided by the function. Function names should be prefixed with the name of the module to which they belong. “Obvious” names (e.g. “util”) for modules should be avoided. Function names should be kept short, if possible. Overloading of function names (using the same function name, differentiating only by input/output parameters, shall be avoided to ensure intent versus determinism.

### Guidelines

Function names should normally be formed from two parts: an action (verb) and an object (noun) of the action. Each word forming the function name is capitalized; except for the first which should be lowercase (this will be the module name). Examples of acceptable function names are chatCloseSession, gpsSetTime.

## Parameter Names

### Standards

Names should be formed from composite words, delimited by upper case letters, with no underscores allowed.

All constants and define parameters will be declared in all upper case.

All macros will be declared as all upper case.

Example:

MYMACRO

## Other Names

### Standards

Enumeration constants and global constant and typedef names should  be stylistically distinguished from function, and variable names. Types (struct types, etc.) should be given meaningful names and begin with an uppercase letter in order to distinguish them from variables.

Example:

MYCONSTANT

MYDEFINE

### Guidelines

Constant names should be in uppercase, with multiple words separated by an underscore. Valid examples are MAX_COLUMN or THOUSAND.

- Macros, with or without parameters, should also be in all upper case.

- numeration constants and global typedef names should be in all uppercase with individual words separated by underscores, e.g. DATA_VALID.

## Variables, Operators and Expression

### Standards

All type casting will be explicit when converting types.

Pointer assignment shall be between pointers of the same type (void * is an exception, though should be used only with good reason). Safe / Static type casting only.

To check the validity of a pointer, compare it to a typecast NULL or integer 0.

### Guidelines

To improve clarity use parentheses even when not required. Using parentheses also helps avoid errors caused by misunderstood operator precedence.

## Functions/ Methods

### Standards

The ANSI-C style will be used for function declaration, with parameter types in the function declaration and a return type declared.

Example:

int power (int base, int n)

Assumptions should not be made about the order of parameter evaluation.  Always explicitly code the desired order of evaluation of sub-expressions.

Example:

func(x * i, y / i, i++)

may be evaluated from right to left or left to right with different compilers and different CPUs.

When unusual data types are mixed across function boundaries use explicit type casts to make it clear that a type conversion has occurred.

If no return value is required for a function, it should be declared as a void function. If a function has an empty argument list this should be clearly indicated using (void), even though many compilers accept ().

Example:

void myVoidFunction(void)

return (void)

## Statements and Control Flow

### Standards

A default selection shall always be provided for switch statements.

Always use braces to delimit the block in an “if” statement, even if there is no else block. A complete if (condition) statement that fits on one line is an exception.

Always use braces to delimit the block in while and do … while statements.

No “goto” statement shall be used.

### Guidelines

All loop structures should have a loop termination state that is used to stop the loop.  The “For” loop uses an increment or decrement counter at the top of the structure. Additional increment or decrement of the loop counter inside the loop should be avoided.

In a “while” loop, the loop termination state should be at the top or bottom of the block with logic updating the loop termination state anywhere in the structure, though counters should be limited in the same fashion as done with the “for” loop.

Break each case of a switch statement to avoid falling through to the next case. The only exception is multiple case labels.

## Code Commenting

### Standards

All files, both header and source, must begin with the standard header information, which contains the file identification information, a one line description of the module, implementation details and indirect reference to the applicable requirements, and the copyright notice.

The header may also contain a longer description of the purpose of the module and any other pertinent information about the module. At a minimum the module name and change history must be including.

A change history for each module must be maintained in a manner appropriate to the development environment.

Example:

/* Author		PCR			Change Description			*/

/* Jon Palmer		[TBD SW ACRONYM]_PCR_0001	Initial development	*/

Block style and in-line comments are both acceptable.

Block style comments should be preceded by an empty line and has the same indentation as the section of code to which they refer. Block style comments should appear at the beginning of the relevant segment of code.

Example

Block Style:

/*

...............

....................

.......................

*/

Brief comments on the same line as the statement that they describe are appropriate for the in-line commenting style.

Example:

In-Line Comments:

...........; /* ......... */

...........; // .........

...........; // .........

Use in-line comments to document variable usage and other small comments.  Block style comments are preferable for describing computation processes and program flow.

### Guidelines

In-line or block commenting rules will be subject to what is supported by the compiler. Most compilers extend the C language to support the In-line (//) comment even though the source module is not a C++ file. If that is not the case, then only block commenting is allowed.

## Header Files

### Standards

Header files should be used as a means of interface specification for a source module. For example, for a module motion_control.c, the external interface containing data declarations, function prototypes and all information necessary to use the module should be declared in motion_control.h.  Requirement identifiers or text should not be included within headers (or code comments), to ensure that such requirements solely reside within the applicable requirements and traceability documents; this will prevent dual maintenance due to multiple locality.

For each function prototype in the header file, the following information should be given in comments:

What the parameters are, how they are used, and any preconditions that must

be established for the parameters.

What the function is going to do to the parameters.

What the results/return values are in all the different cases possible for the function.

To avoid nested includes, use #define wrappers as follows:

#ifndef MOTION_CONTROL_H

#define MOTION_CONTROL_H

......

......

#endif /* MOTION_CONTROL_H */

## Source Files

### Standards

Code comments should be used to give an English language synopsis of a section of code, to outline steps of an algorithm, or to clarify a piece of code when it is not immediately obvious what was done or why it was done. In no case should the code comments just parrot the code.

Use consistent nested indentation throughout your code. When working in existing code, respect the indentation style currently in use. At a minimum, indentation should be consistent within a function.

### Guidelines

Code comments which apply to a block of code (either a loop or branch construct, or a grouping of statements), should be immediately above the block and indented to the same level as the code.

Code comments which apply to a single statement may be immediately above or to the right of the statement.

## Readability

### Standards

Write only one statement per line; “elegant” source code construction combining constructs that could be more readily placed upon multiple lines should be avoided.  Remember, source code is only as good as the weakest link, meaning the least experienced or capable individual constructing, reviewing, or testing that code must be able to meet the applicable safety criteria.

### Guidelines

Use blanks around all binary operators except “.” and “->”.

In compound statements, put braces ({}) in a separate line and aligned with parent statements. Example:

{

......;

......;

}

Use a blank after commas (arguments list, values, etc), colon, semicolon and control flow keywords:

procedure (arg1, arg2);

for (i = 0, j = strlen(s) – 1; i < j; i++, j--;)

if (...)

{

…..;

}

while (...)

{

…..;

}

Do use blanks between an identifier and any of “(“, “)”, “[“ or “]”, with the exception of two or more “(“, “)”, “[“, or “]”:

procedure ( arg1, arg2 )

a[ I ] = 1;

z[ I ] = a + ( sin ( x ) + cos ( b [ i ] ));

printf ( “%d”, ( a + b ));

Line up continuation lines with the part of the preceding line they continue:

a = ( b + c ) *

( c + d )

if (( a == b ) &&

( c == d ))

printf ( “%d %s”,

( a + b ),

“abcd” );

Indents should be spaces instead of tabs.

Object declarations should be visually aligned as it is easier to scan a list of objects when names and types appear in the same column. However, this guideline should be used with some common sense.

Examples:

struct FL_ENTRY *	fl_p;

struct FLDSYM *	sym_p;

char * 	wsbuf;

int 	length;

The following templates provide examples of how to indent the most commonly used code structures. Note that braces on a new line are the style here, but braces on the same line as the opening statement are also acceptable, but remember that consistency is always a goal.

typedef struct

{

...... ......;

...... ......;

...... ......;

} ......;

if (....)

{

......;

......;

}

else

{

......;

......;

}

if (....)

{

......;

......;

}

else if (....)

{

......;

......;

}

else if (....)

{

......;

......;

}

switch (....)

{

case ...:

......;

......;

break;

case ...:

......;

......;

break;

.

default:

......;

......;

break;

}

while (....)

{

.....;

......;

}

for (...; ...; ...)

{

......;

......;

}

do

{

......;

......;

}

while (....);

# Code Reviews

This standard may be used when performing code reviews of final firmware/software developed for EPIC’s quality system and/or products.  Code reviews can consist of the following activities:

## Code Inspection- A manual [formal] testing [error detection] technique where the programmer reads source code, statement by statement, to a group who ask questions analyzing the program logic, analyzing the code with respect to a checklist of historically common programming errors, and analyzing its compliance with coding standards. Contrast with code audit, code review, code walkthrough. This technique can also be applied to other software and configuration items. Syn: Fagan Inspection. See: static analysis.

## Code Review- A meeting at which software code is presented to project personnel, managers, users, customers, or other interested parties for comment or approval. Contrast with code audit, code inspection, code walkthrough. See: static analysis.

## Code Walkthrough- A manual testing [error detection] technique where program [source code] logic [structure] is traced manually [mentally] by a group with a small set of test cases, while the state of program variables is manually monitored, to analyze the programmer's logic and assumptions. Contrast with code audit, code inspection, code review. See: static analysis.

Document Revision History

| Change Control Number: | Version Number: | Description of Change: | Date: | Updated by: |
| --- | --- | --- | --- | --- |
| CC-120228-01 | 000 | Introduction | 2/28/12 | A. Mason |
| CC-130319-02 | Obs | Obsolete document. | 3/21/13 | L. Kurth |