- Purpose

The purpose of this procedure is to provide instruction for configuration of the USB Hardware Key(s) for use with Computer Systems (1282-1511) loaded with the ClearView Software.

- Scope

This procedure is intended to be used to configure any USB Hardware Key(s) whether used in development, clinical applications, and/or in saleable product.

- Responsibilities

| CTO | Responsible for overseeing the configuration process and ensuring all process steps and documentation requirements are followed. |
| --- | --- |
| Purchasing | Responsible for the acquisition of USB Hardware Key(s). |
| Software Installer | Responsible for following this procedure and generating all required documentation. |

- Equipment

Manufacturer’s HASP Development Kit

USB Hardware Key, 1282-1515

USB Hardware Key Label, 1282-1523

USB Hardware Key Configuration Record, ITF-008

Zip lock bag

- HASP Envelope setup:

  - Obtain a USB Hardware Key Configuration Record (Record) and USB Hardware Keys (1282-1515) to be configured.

  - Record the revision level of this procedure in the space provided on the Record.

  - Record the revision level of the USB Hardware Keys provided on the Record.  If more than one revision is received from inventory, create separate Records for each revision received.  If an EPIC lot number has been assigned, record the EPIC lot number in the space provided on the Record.

  - Record the manufacturer’s configuration software name and version in the space provided on the Record.

  - Record the serial number (i.e., unique series of numbers/letters located on the back of the USB Hardware Key) for each USB Hardware Key in the column provided on the Record.

  - Perform the following configuration steps for each USB Hardware Key.  If the configuration is completed correctly, record Pass in the Pass/Fail column of the Record.  If the configuration is not completed correctly, record Fail in the Pass/Fail column of the Record.

    - Install the HASP Development Kit.

    - Introduce master key by running MasterHASP Wizard.

    - Secure ClearView using the Envelope software from SafeNet.

    - Drag and drop the final compiled ClearView executable in to the Envelope window.

    - Set the User Message for error #7 to “ClearView key not found”.

    - Select “Protect” from the ClearView executable configuration page under “Programs”.

    - A new ClearView executable is generated with the protected code.  This executable should be released with the rest of the system.  Copy the generated .exe file in to the installation folder, replacing the ClearView.Client.exe; this will require the HASP key in order to run the software.

  - If USB Hardware Key(s) have failed the configuration process, separate these items from those that pass, generate a Non-conforming Material Report (NCMR), and label product as “Rejected” (including the part number, revision level, lot number (if applicable), serial number, and the label “Rejected”).  Obtain a NCMR number from QA and record the number on the NCMR and the Record.  Forward the keys and the NCMR to QA for quarantine and disposition per the Non-conforming Material procedure.

  - Each USB Hardware Key that passes the configuration management requirements will be placed into a zip-lock bag and labeled with a USB Hardware Key Label (1282-1523).

  - Complete the configuration documentation by signing and dating the Record in the space provided.  Print your name and record your department in the space provided on the Record.  Forward all packaged USB Hardware Keys and the Record to QA for inspection and release.

  - QA Inspection and Release- For each of the following inspection points, record Pass or Fail in the appropriate column on the Record.  If any inspection fails, record Fail in the Pass/Fail column of the Record.

    - Verify that the Record has been completed accurately and that the record keeping meets all quality system requirements.

    - Verify that each USB Hardware Key is packaged separately

    - Verify that each USB Hardware Key is labeled with the current version of the USB Hardware Key Label (1282-1523).

    - Place all keys that meet that pass all inspections in a container.  Label the container with a green Approved label.  The labels must clearly state “Approved” and contain the product drawing/specification number, revision level, quantity, lot number (if applicable), and/or acceptance date.  Initial and date each label after verification of the accuracy of the information recorded.  Place a label on each container of keys.

    - Place all keys that do not meet all inspection in a container.  Generate a Non-conforming Material Report (NCMR), and label product clearly stating “Rejected” and including the product drawing/specification number, revision level, quantity, lot number (if applicable), and/or rejection date.  Obtain a NCMR number and record the number on the NCMR and the Record.  Forward the keys and the NCMR to QA for quarantine and disposition per the Non-conforming Material procedure.

    - Complete the inspection documentation by signing and dating the Record in the space provided.  Print your name and record your department in the space provided on the Record.

    - Forward all acceptable keys to inventory control.

- References

1282-1515, USB Hardware Key

1282-1523, USB Hardware Key Label

ITF-008, USB Hardware Key Configuration Record

QA-017, Non-conforming Material

QAF-011, Non-conforming Material Report

Document Revision History

| Change Order: Number: | Version Number: | Description of Change: | Date: | Updated by: |
| --- | --- | --- | --- | --- |
| CC-120517-01 | 000 | New Document | 5/17/12 | A. Mason |