| Procedure: IT-011 | Revision Level: |
| --- | --- |
| Computer System Part Number: 1282-1511 | Revision Level: |
| Computer System Lot Number (if applicable): |  |
| Computer System Manufacturer’s Serial Number: |  |
| EPIC Internal Computer System Reference Number: |  |
| USB Hardware Key Part Number: 1282-1515 | Revision Level: |
| USB Hardware Key Lot Number (if applicable) |  |

| Installation Verification |  |  |
| --- | --- | --- |
| Steps | Complete (Initial) |
| Verify computer system has Windows XP 32bit software running as the operating system. If not, install. |  |
| Apply all relevant Microsoft patches, using the Microsoft Update feature. Also verify that Automatic Update is set to ‘Download updates for me, but let me choose when to install them’. |  |
| Change the way users log in to not use fast switching nor the welcome screen. |  |
| Set the proper passwords for ‘EPICAdministrator’ and ‘EPICUser’ |  |
| Log out as Administrator and log back in as EPICAdministrator. |  |
| Install the Full .Net 4.0 Framework, and patches. |  |
| Load the necessary printer driver for the system. |  |
| Install the Camera Drivers. |  |
| Install the latest FTDI Drivers. |  |
| Assign the proper FTDI Driver. |  |
| Install the database engine. |  |
| Create and install the blank ClearView database, remember to set the MAC address in the CaptureDevice table. |  |
| ClearView Software Version Number: |  |
| Install the HASP run-time engine. |  |
| Install the Crystal Reports Run-time Engine files. |  |
| Install the Matlab Run-time Engine. |  |
| Install the ClearView Shell application. |  |
| ClearView Shell Version Number: |  |
| Create the folder “ClearView” in “Programs Files\EPIC” on 32bit systems or “Program Files (x86)\EPIC“on 64bit systems. |  |
| Place the files from the ‘Release’ folder of the version that is being deployed into the folder created in step 5.17. |  |
| Create shortcut on desktop named “ClearView” which points to executable “EPIC.ClientDeviceUI.exe” in ClearView folder. |  |
| Copy the appropriate license file from the ‘licenses’ folder to the ClearView folder. |  |
| ClearView License: |  |
| Modify the config file to have the proper parameters |  |
| Insert a hasp key and verify the application starts up and a login screen is displayed. |  |

Installation Failure:

Notes:

| Name: | Department: |
| --- | --- |
| Signature: | Date: |
| Operational Verification |  |
| ClearView Device Serial Number: |  |

| Steps | Pass/Fail Result (circle one) |  | Complete (Initial) |
| --- | --- | --- | --- |
| Reboot system to establish a baseline starting point. | Pass | Fail |  |
| Log in as EPICAdministrator using password 567dop923rap7#. | Pass | Fail |  |
| Double click on ClearView icon. Login screen displays. | Pass | Fail |  |
| Log in as user “TechAdministrator” with password 473pol965Drp#. | Pass | Fail |  |
| Log in as user “Administrator” with password 329dap735r#. | Pass | Fail |  |
| Create the EPICAdministrator user using the following information: Username: EPICAdministrator Password: c13ArV123# First Name: EPIC Last Name: Administrator Email: epicadministrator@epicdiagnostics.com Role Assignment: CustomerAdmin, CustomerUser | Pass | Fail |  |
| Create the EPICUser user using the following information: Username: EPICUser Password: c13ArV13W# First Name: EPIC Last Name: User Email: epicuser@epicdiagnostics.com Role Assignment: CustomerUser | Pass | Fail |  |
| Look at bottom of application and verify the Camera Status and Scanner Status are “Available”. If they are not, try disconnecting and reconnecting USB cable a few times before taking corrective action. | Pass | Fail |  |
| Add a new patient named “Application T. Test” using any gender and birthdate. Make sure the user can be saved (created). | Pass | Fail |  |
| Verify that a successful calibration process can be performed. | Pass | Fail |  |
| Perform a patient capture using the patient Application T. Test with the calibration probe and verify that the scanner captures images and can perform a complete calculation process. | Pass | Fail |  |
| Log out of the application. | Pass | Fail |  |
| Log into the application using the EPICUser login | Pass | Fail |  |
| Perform a calibration logged in as EPIC user and verify that the scanner captures calibration images. | Pass | Fail |  |
| Perform a patient capture logged in as EPICuser using the patient Application T. Test with the calibration probe and verify that the scanner captures images and can perform a complete calculation process. | Pass | Fail |  |

Operational Failure:

Notes:

| Name: | Department: |
| --- | --- |
| Signature: | Date: |

Labeling Information

Computer System with ClearView Software Box Label Part number: 1282-1520 Revision level:

ClearView Version recorded on label:

| QA Inspection |  |  |
| --- | --- | --- |
| Inspection Point | Pass/Fail | Comments/Notes |
| Verify that this record is completed correctly and accurately |  |  |
| Verify that computer is labeled with EPIC reference number (if applicable) |  |  |
| Verify that the computer is labeled with the version number of the ClearView software installed |  |  |

| Inspected By: |  |
| --- | --- |
| Signature: | Date: |
| Print Name: | Department: |

Document Revision History

| Change Order: Number: | Version Number: | Description of Change: | Date: | Updated by: |
| --- | --- | --- | --- | --- |
| CC-120517-01 | 000 | New Document | 5/17/12 | A. Mason |