- Purpose

To establish procedures for the inspection and release of ClearView Device, Packaged (EPC-100) received from EPIC’s contract manufacturer.

- Scope

This procedure applies to all packaged ClearView Devices received from EPIC’s contract manufacturer.  This procedure is used by Quality Assurance personnel that are responsible for product quality.

- Definitions

| Device History Record (DHR) | A compilation of records containing the production history of a finished device. |
| --- | --- |
| Device Master Record (DMR) | A compilation of records containing the procedures and specifications for a finished device. |
| Finished Device | A device or accessory to any device that is suitable for use or capable of functioning, whether or not it is packaged, labeled, or sterilized (if applicable). |
| Non-conformity | The nonfulfillment of a specified requirement. |

- Responsibilities

| User | - Complies with the policy and procedure. - Completes all required documentation. - Ensures the most current version of this document is used when referenced. |
| --- | --- |
| Departmental Management | - Ensures departmental personnel are properly trained before using this policy or procedure. - Ensures a timely inspection and approval/rejection of materials. |
| Executive Management | - Ensures the requirements of this plan are understood and implemented throughout all applicable levels of the company. |
| Quality Assurance | - Complies and enforces the standards within this document. - Monitors the implementation and effectiveness of this document. - Audits to ensure compliance with the referenced procedures. |

- Procedure

EPIC Research & Diagnostics, Inc. (EPIC) utilizes a contract manufacturer for the assembly and test of the ClearView Device.  The ClearView Device is defined in product specification, EPC-100 and the following steps will be completed prior to release of these devices for further inspection and test.

  - Receiving Inspection

    - Materials are received and handled according to the Receiving procedure.  The Receiving Clerk will document the receipt of ClearView Devices on the Receiving Log, label the containers and forward the material to QA for quarantine until completion of the inspection process.

    - Receiving inspection is performed on all ClearView Devices manufactured by the contract manufacturer and documented on the Receiving Inspection Record.   This inspection is intended to review the documentation and not perform any device performance verification.  Complete the information at the top of the ClearView Device, Packaged Receiving Inspection Record (Record).

    - The QC Inspector will locate a copy of the PO, the ClearView Device, Packaged, EPC-100 product specification, Packing List and the Device History Record (DHR) for the device under inspection.

    - All records must be legible and document corrections made must meet the requirements established in EPIC’s Document Corrections procedure.  Any information htat is not legible and/or corrections that are made that do not meet the requirements of EPIC’s Document Corrections procedure will be considered a non-conformity and result in rejection of the device for disposition by the Material Review Board (MRB) according to EPIC’s Non-conforming Material procedure.

    - All records must be complete and accurate to the requirements.  Perform the following document verifications:

| Document | Verification | Acceptance Criteria |
| --- | --- | --- |
| Packing List | Verify that the Customer Order No. matches EPIC’s PO number. | Records must match |
|  | Verify that the part number and revision level match EPIC’s PO. | Records must match |
|  | Verify that the serial number recorded on the Packing List and the device packaging label on the product shipping box. | Records must match |
| Certificate of Compliance (C of C) | Verify that the P. O. Number matches EPIC’s PO number. | Records must match |
|  | Verify that the part number and revision level match EPIC’s PO, the Packing List, and EPIC’s current released part numbers. Review Document Control revision history to confirm if any rework/processing instructions exists for previous revisions. Follow approved instructions if present. | Records must match |
|  | Verify that the serial number recorded on the Packing List and the device packaging label on the product shipping box. | Records must match |
|  | Verify that the C of C has a statement certifying that the material and parts are being certified as being manufactured and inspected according to all referenced procedures. | Statement must be present on the C of C |
|  | Verify that the C of C provides the contract manufacturer’s name, address, and contact information. | Name, address, and contact information must be accurate |
|  | Verify that the C of C has been signed and dated by an appropriate representative of the contract manufacturer. | - Must have an original signature of employee with appropriate authority to provide the certification (manager level and above) - Must be dated after all other DHR documents (i.e., should have been completed after all other records with the DHR) |
|  | Record any contract manufacturer referenced non-conformances and/or deviations in the space provided on the Record. | N/A |
| Device Traceability Record(s) | Verify that the product part number and revision level match the DHR records. | Records must match |
|  | Verify that the contract manufacturer traveler ID matches the DHR records. | Records must match |
|  | Verify all traceability information is recorded (no blank data entry spaces), and each traceability is initialed and dated by the technician and the QC Release. | Record must be complete and accurate. |
|  | Verify that serial numbers are recorded for the following part numbers listed on the traceability record(s): |  |
|  | 06690, ClearView Assembly Build | Serial number must be recorded |
|  | MPS-65-48, Power Supply | Serial number must be recorded |
|  | 06686, Power entry Module Assembly | Serial number must be recorded |
|  | 1282-1016, Conductive Glass Assembly | Serial number must be recorded |
|  | Flood LED PCBA | Serial number must be recorded |
|  | Verify that lot numbers are recorded for the following part numbers listed on the traceability record(s): |  |
|  | 06686, Power entry Module Assembly |  |
|  | All cables including part numbers: 06670, 06624, 06682, 06681, 06636, 06673, and 06684 |  |
|  | SPC900NC/27 ccd Camera |  |

| Document | Verification | Acceptance Criteria |
| --- | --- | --- |
| Assembly Records | Obtain a copy of each of the following assembly procedures. Record the revision of each procedure used during inspection in the space provided on the Record. Verify the DHR documentation required by each has been generated as required by the approved procedure. If a procedure is not performed, provide the deviation and/or non-conformance reference number authorizing the change to the process. |  |
|  | XEPC-01_1 ClearView (XEPC01A) Cover Housing Assy WI |  |
|  | XEPC-01_2 ClearView (XEPC01A) Base Housing Assy WI |  |
|  | XEPC-01_3 ClearView XEPC01A Final Assy and Acceptance Testing |  |
|  | XEPC-02 Electrode Assy WI |  |
|  | XEPC-03 Electrode Board WI |  |
|  | XEPC-04 Flood LED PCBA WI |  |
|  | XEPC-05 Control Board PCBA WI |  |
|  | XEPC-06 Power Entry Module WI |  |
|  | XEPC-07 Camera Cable WI |  |
|  | XEPC-09 DC Power Cable WI |  |
|  | XEPC-10 AC Power Cable WI |  |
|  | XEPC-11 Flood LED Cable WI |  |
|  | XEPC-12 Rear Panel LED WI |  |
|  | XEPC-13 Shield to Ground Cable WI |  |
|  | XEPC-14 USB Panel Mount Cable WI |  |
| Test Records |  |  |
| Nonconformance Reports (if applicable): | Non-conformance reports are generated wherever appropriate throughout the manufacturing process. If a non-conforming report is listed on a document being reviewed, record the reference number on the Record. Perform the following verifications: |  |
|  | Verify that the report has a complete description of the problem. | Record must be accurate and complete. |
|  | Verify that the investigation and evaluation of the non-conformance is documented on the report. | Record must be accurate and complete. |
|  | Verify that disposition of the non-conforming product has been documented on the report. | Record must be accurate and complete. |
|  | If disposition included rework, verify that the rework instructions are written and approved by the same disposition authority and includes a representative from EPIC. | Record must be accurate and complete. |
|  | If disposition is use as is (or similar), verify that disposition approval includes a representative from EPIC. | Record must be accurate and complete. |
| Deviations (if applicable, list deviation numbers) | Deviations are generated wherever appropriate throughout the manufacturing process. If a deviation is listed on a document being reviewed, record the reference number on the Record. Perform the following verifications: |  |
|  | Verify that the deviation has a complete description of the deviation from approved procedure. | Record must be accurate and complete. |
|  | Verify the deviation approval includes a representative from EPIC. | Record must be accurate and complete. |
|  | Verify that the deviation instructions regarding recordkeeping have been met in the DHR. | Record must be accurate and complete. |
| Final QC DHR Review Records |  |  |

    - If the paperwork verification points all meet the release criteria, mark the Pass on the Receiving Inspection Record.  If any inspection point does not meet the release criteria, mark Fail on the Receiving Inspection Record.  Sign and date in the space provided.

    - ClearView Devices are currently in development.  Devices are not released for use from this procedure, but are released for further performance verification and validation testing conducted under EPIC’s Design Control Procedure.  Forward materials to the appropriate stock location.  File the inspection report by product drawing/specification number and received date.  Inspection records are to be maintained in secure location with limited access.

    - For components/materials that do not meet the release criteria, generate a Non-conforming Material Report and fully document the cause of the non-conformance.  Generate package labels for each container.  The labels must clearly state “Rejected” and contain the product drawing/specification number, revision level, quantity, lot number (if applicable), and/or received date.  Initial and date each label after verification of the accuracy of the information recorded.  Place a label on each container of components/materials.

    - Forward materials to the Quarantine location.  Notify the Quality Assurance Manager of the failure.  File the inspection report by product drawing/specification number and received date.  Inspection records are to be maintained in secure location with limited access.

  - In-Process Inspection- This section is reserved for in-process inspection instructions that will be developed prior to the need for in-process inspection.

  - Final Inspection- This section is reserved for final-process inspection instructions that will be developed prior to the need for in-process inspection.

- Reference Documents

EG-001, Design Control Procedure

OPS-004, Receiving Process

OPS-006, Purchasing

QA-007, Document Corrections

QA-017, Non-Conforming Material Process

QAF-011, Non-Conforming Material Report

QAF-015, ClearView Device, Packaged Receiving Inspection Record

- Attachments

None.

Document Revision History

| Change Order Number: | Version Number: | Description of Change: | Date: | Updated by: |
| --- | --- | --- | --- | --- |
|  | 000 | New Document |  | A. Miller |