- Purpose

This procedure establishes the methods to identify product during all stages of assembly, manufacture, testing, and delivery.

- Scope

This procedure applies to all EPIC Research and Diagnostics, Inc. (EPIC) personnel and supervision that handle product throughout all phases of assembly, manufacture, testing, and delivery.

- Definitions

| Identification | Labeling of the component/material in order to distinguish one item/batch/trace/lot from another. |
| --- | --- |
| Nonproductive Materials and Services | Items not used in the actual production of devices and/or support products. Nonproductive materials include office supplies, cleaning products (not associated with device manufacture or maintenance), etc. |
| Productive Materials | Materials used in the actual production of devices and support products. Productive materials are only purchased from approved suppliers. |
| Traceability | Labeling and handling of a component/material in a way that identifies the history (receipt, part number/version, etc.) of the item. |

- Responsibilities

| User | - Complies with the policy and procedure. - Ensures the most current version of this document is used when referenced. - Maintains all documentation required to ensure traceability of all materials and products. - Labels all materials and product as required by EPIC’s quality system policy and procedures. |
| --- | --- |
| Departmental Management | - Ensures departmental personnel are properly trained before using this policy or procedure. - Ensures materials inventory is maintained in a manner that ensures all product is constantly traceable to the requirements specified by product specifications and traceable as required. |
| Quality Assurance | - Monitors the implementation and effectiveness of this document. - Conducts periodic audits to assure compliance. |

- Policy

  - Identification

EPIC establishes and maintains procedures for the identification of components, assemblies, and finished product as appropriate for their purpose, use and life cycle.

EPIC ensures that products are controlled through the use of alphanumeric part numbers and/or control numbers.  Additional identification may be recorded such as serial, batch, and/or trace numbers, as applicable.  All incoming, in-process and/or finished device identification requirements will be defined with instructions for execution of this identification throughout processing.  Product numbers and control numbers are recorded on appropriate documentation throughout the process as required by standard operating procedures and all records are filed accordingly.

EPIC ensures that all contract manufacturing sites have implemented procedures to ensure the identification of all components, assemblies, and finished product that are handled throughout their specific contract manufacturing site quality system in accordance with the requirements specified in the product specifications and process procedures.

  - Traceability

EPIC establishes and maintains procedures for the traceability of components, assemblies, and finished product, and ensures that they are controlled through the use of alphanumeric part numbers/control numbers, such as serial, batch, or trace numbers and date received, as applicable.  Product numbers and control numbers are recorded on appropriate documentation throughout the process as required by standard operating procedures and all records are filed accordingly.

- Procedure

  - Productive material’s identification and traceability begin upon arrival at EPIC facilities.  The procedure for meeting the identification and traceability requirements upon receipt is provided in the Receipt and Storage of Materials and Finished Devices.

  - All contract manufacturing facilities will be required to have procedures that comply with regulatory requirements regarding identification and traceability.  EPIC will verify this compliance through the Internal Audit procedure and/or periodic review of product quality at Receiving Inspection.

  - Upon delivery, all materials and components are held in a clearly identified quarantine area where they are reviewed for their identification and traceability requirements.  The product/component specification and this procedure will be used to identify these requirements.  All received items not delivered immediately to the Purchase Order Requestor and those being held as raw material, component, and/or finished device inventory are to be properly labeled prior to release from quarantine.

  - Non-productive materials and components which cannot be immediately delivered to the recipient are required to be labeled with the following information.  This information can be in the form of a hand-written or type-written label or simply handwritten on the container.  This labeling will be easily identifiable at all times.

Non-Productive Material

Traceability Not Required

Received Date: ___________

Qty: ___________________

  - Productive materials and components are required to be labeled according to their traceability requirements as well as identified by their proper status (i.e., Approved, Unapproved, and Rejected).  The traceability requirement varies depending upon the application of the material or component.  The Part Number Traceability Matrix provides a part number specific list of items and identifies their traceability requirement.

    - Serial Number Traceable – These materials/components are required to have traceability for each and every unique item.  A unique alpha/numeric numbering system will be established to identify device serial numbers.  The number assigned must be historically unique for each part number (no duplicates).  Each material/component must be identified as to the unit of measure as well as the serial number for that material/component.  Each unit will then be required to be labeled with the product name and/or number as well as the unique serial number for the unit.  The labeling may exist as a single label or may be constituted with several labels.

    - Lot Number Traceable – These materials/components are required to be traced throughout assembly and distribution and can be grouped into batches or lots.

      - The traceability is established by labeling the product upon receipt. If productive material has not been assigned a lot number by the supplier, a trace number is assigned comprised of the date received followed by a unique 2-digit identifier.

      - Each set/batch/box of materials/components must be identified with the following information:

Or appropriate approval status

UNAPPROVED

Part Number:_______________  Rev.: _______

Name: _________________________________

Lot Number: ___________

Qty: ____________________

    - Traceability Not Required – These materials/components are not required to be lot number or serial number traceable.  Identification is maintained by labeling each container as follows: This information can be in the form of a hand-written or type-written label or simply handwritten on the container.  This labeling will be easily identifiable at all times.  If multiple shipments are received on the same day a sequential identifier “-01”, “-02”, etc. will be added to the date received and recorded on the packing list.

Or appropriate approval status

UNAPPROVED

Traceability Not Required

Part Number:_______________  Rev.: _______

Name: _________________________________

Date Received: __________________________

Qty: ____________________

  - Approval Status – For productive products, the status of the materials/components must also be identified.  A color-coded system will be established to ensure that the status of the product is easily identifiable at all times.

    - Approved – Materials and/or components that have met all release criteria are able to be marked Approved.  This approval indicates that the material and/or component may be released for use as established in the product specification and/or the manufacturing policies and procedures.  The word “Approved” will be clearly visible on any and all labeling.  These labels do not all have to be completely green, but should have a portion visibly green in color and easily identifiable as approved.

    - Unapproved – Materials and/or components that have not been reviewed to ensure quality system requirements have been met.  These items are required to be stored in a quarantine area separate from the assembly process to prevent mix-up between released “Approved” items and “Unapproved” items.  These labels do not all have to be completely yellow, but should have a portion visibly yellow in color and easily identifiable as unapproved.

    - Rejected – Materials and/or components that have been reviewed to ensure that the quality system requirements have been met and do not meet the release criteria must be labeled “Rejected” and set apart from approved and/or unapproved items.  These labels do not all have to be completely red, but should have a portion visibly red in color and easily identifiable as rejected.

- Reference

OPS-004, Receipt and Storage of Materials and Finished Devices

OPSF-005, Part Number Traceability Matrix

- Attachments

N/A

Document Revision History

| Change Order Number: | Version Number: | Description of Change: | Date: | Updated by: |
| --- | --- | --- | --- | --- |
| CC-100528-02 | 000 | New Document | 05/28/10 | A.M.M. |
| CC-130215-01 | 001 | Revisions to clarify lot number assignments and make minor punctuation edits | 2/19/13 | J. Ayres |