- Purpose

To establish procedures for ensuring that products that do not conform to specified requirements are prevented from unintended use.

- Scope

This procedure applies to all EPIC personnel and supervision that are responsible for product quality.  EPIC Quality Assurance management is responsible to integrate the requirements of this policy and procedure with any contract manufacturer’s non-conforming material procedure.

- Definitions

| Corrective Action | The action(s) taken to correct a non-conformance or make materials suitable for use. |
| --- | --- |
| Deviation | Unauthorized modification of an approved process. |
| Disposition | Instructions or actions required to control the use of non-conforming material. |
| Non-conformance | Materials and/or product that does not meet requirements and/or specifications. |
| Preventive Action | The action(s) required to prevent a non-conformance from happening again. In the case of material disposition, preventive action is suggested and actual implementation is analyzed and documented through the Corrective and Preventive Action Plan. |

- Responsibilities

| User | - Document any discrepancy and/or non-conformance (whether in process and/or product) identified during the performance of quality system procedures. - Marks all non-conforming materials as required an segregates any non-conforming materials - Complies with the policy and procedure for all identified discrepancies/non-conformances. - Ensures the most current version of this document is used when referenced. |
| --- | --- |
| Departmental Management | - Ensures departmental personnel are properly trained before using this policy or procedure. - Ensures that all responsibilities within their assigned department are completed as required in this procedure. |
| Quality Assurance | - Manage and oversee the Material Review Board (MRB) and all trend tracking associated with material review. - Receives all non-conformance documentation and product and processes the product and documentation through the material review process. - Monitors the implementation and effectiveness of this document. - Management representative acts as the facilitator of the MRB, ensuring that all documentation from the MRB meets EPIC’s quality system requirements. - Conducts periodic audits to assure compliance to this procedure. |
| Manufacturing | - Management representative acts as a member of the MRB as required by policy and requested by the MRB facilitator. - Investigates, evaluates and responds to all assigned requests for information related to discrepancies/non-conformances. - Responsible to execute any and all assigned actions as required by the MRB. |
| Engineering | - Management representative acts as a member of the MRB as required by policy and requested by the MRB facilitator. - Investigates, evaluates and responds to all assigned requests for information related to discrepancies/non-conformances. - Responsible to execute any and all assigned actions as required by the MRB. |
| Marketing | - Management representative acts as a member of the MRB as required by policy and requested by the MRB facilitator. - Investigates, evaluates and responds to all assigned requests for information related to discrepancies/non-conformances. - Responsible to execute any and all assigned actions as required by the MRB. |

- Policy

EPIC establishes and maintains procedures to ensure components, assemblies and finished devices that do not conform to the quality system requirements at any point throughout the process are identified and segregated until disposition and the person(s) responsible notified accordingly.  Records of non-conformances are maintained and reviewed by designated individual(s).  Components, assemblies, and/or finished devices will be dispositioned by the appropriate personnel.  Any proposed rework is performed to established procedures, reevaluated and retested. All non-conformances related to product will be evaluated for appropriate corrective and preventive action.  These records are a part of the Device History Record.

- Procedure

A process flowchart, EPIC’s Non-Conforming Material Process Flowchart, has been created to help facilitate understanding of the non-conforming material process.

  - Identification of Process Discrepancies and/or Non-conforming Product

For the purposes of this process, discrepancies in the execution of processes associated with materials processing and/or product assembly, test, and inspection are considered “non-conformances” and will be handled as a non-conformance and materials/product associated will be processed as non-conforming materials.

    - Non-conforming product can be identified by any EPIC employee at any point within the process.

    - The materials and/or product affected are immediately removed from the normal process flow.

    - The materials must be identified with a hold tag or label.  The label must clearly identify the material as being on hold and indicate the part number, quantity, description, reason for the hold, name of the person who detected the problem and the date.

    - The individual who identify the non-conformance is required to document the non-conformance on the Non-Conforming Material Record (NCMR). Complete the NCMR by recording all information in Section 1 of the form.

    - A NCMR number is obtained from QA and recorded at the top of the NCMR.

    - QA is responsible for maintenance of the Non-Conforming Materials Report Log and issuing new NCMR numbers.  The QA designee is responsible to enter all relevant information into the Non-Conforming Materials Report Log at the time of NCMR number being issued.

    - The non-conforming material with proper labeling is forwarded to a quarantine location along with the original NCMR and held for the MRB review.

  - Materials Hold

    - Material on hold must be segregated from accepted product, properly labeled as to the status and prevented from use until MRB disposition.

    - Material on hold cannot be removed from hold without prior authorization through an MRB disposition or process deviation.

    - Materials on hold for longer than 30 days should have a written explanation for the delay attached to the original NCMR.

    - QA is responsible for oversight of the hold materials.

  - Material Review Board (MRB)

    - The Material Review Board consists of appropriate representatives from each department with the technical knowledge and experience to disposition product and suggests corrective and preventive action.  Each department will be responsible to identify one individual who will participate on materials review decisions.

    - The QA Manager or designee will act as a facilitator for the MRB.  The facilitator of the MRB is responsible to notify MRB representatives when materials require disposition.  The facilitator of the MRB is also responsible to oversee the documentation on the NCMR, forward materials and documentation as instructed by the MRB, and process the NCMR for tracking and follow-up after actions are assigned by the MRB.

    - MRB representatives are responsible to represent the technical expertise of their functional group during MRB determinations.  MRB representatives are also responsible to carry out all instructions required by the board within the time frames assigned.

    - MRB reviews the non-conformance and identifies any appropriate actions necessary.  Any instruction for use as is, rework, and/or rejection/scrap will be documented in the Disposition Section.  The department responsible for implementing the disposition should be indicated in the Disposition Section.  An example of possible dispositions include:

- Use as is- Use of the product/component as is without modification or further inspection.  Use as is dispositions require a written justification of the disposition to be included.

- Rework- Rework dispositions require the approval of written rework instructions to be included with the NCMR or handled through EPIC’s Deviations procedure.

- Reject- Reject dispositions are reserved for incoming materials which are being rejected back to the supplier due to the non-conformance.

- Scrap- Scrap dispositions require the proper disposal of the product/components in question.  MRB members who make scrap dispositions must include a representative from the department from which the scrapped items will be charged.

    - The MRB will also assess any corrective action already taken and determine if further corrective action is necessary.  If further corrective action is necessary, the MRB will assess whether a Corrective and Preventive Action Plan (CPAP) or Supplier Corrective and Preventive Action Plan (SCPAP) will be required.  Indicate the need to issue a CPAP or SCPAP by checking the box provided.  If further corrective action is not required, a justification for this decision will be recorded in the Corrective Action section of the NCMR.  The department responsible for implementing the corrective action should be indicated in the Corrective Action Section.

    - The MRB will also assess the need for preventive action.  If further preventive action is necessary, the MRB may suggest the type/method of preventive action that will be implements.  This suggested preventive action will be recorded in the Preventive Action section of the MRB.  If preventive action is required, a Corrective and Preventive Action Plan (CPAP) or Supplier Corrective and Preventive Action Plan (SCPAP) will be required.  Indicate the need to issue a CPAP or SCPAP by checking the box provided.  If further preventive action is not required, a justification for this decision will be recorded in the Preventive Action section of the NCMR.  The department responsible for implementing the preventive action should be indicated in the Preventive Action Section.

    - MRB Review is documented by a signature and date for each required functional group in the appropriate location in Section 3 of the NCMR.  MRB can be conducted with a minimum representation of QA and OPS or ENG.  MRB representatives can request additional representatives at any time as needed to ensure that the proper functional expertise is represented on the board.

  - Assigned departments are required to implement all disposition instructions, corrective action requirements and preventive action requirements.  Product should be properly labeled at all times throughout the process with labeling that clearly identifies the product as non-conforming until the disposition is determined to allow different labeling.  In the absence of clear delineation, QA will take responsibility for ensure that all MRB decisions are carried out.  The Non-Conforming Material Process Flowchart can be used as a reference for proper process flow.

  - NCMR Final Review

Upon completion of all disposition activities, the original NCMR is forwarded to QA for Final Review and filing.  The final review will include the following steps:

    - Review for Completion- The document and any attachments will be reviewed to ensure that all required documentation is completed.  Additionally, the review will ensure that all corrections and additions were made appropriately and attachments are referenced clearly on all documentation.

    - Process Tracking- All required information will be entered into the Non-Conforming Materials Report Log Excel spreadsheet.  The information in the spreadsheet will be used to report on process data relevant to understanding supplier performance and quality system performance.  This information will be reviewed periodically with QA management and executive management.

    - Corrective and Preventive Action Plans- QA will issue all necessary CPAP or SCPAP according to the requirements specified in the Corrective and Preventive Action Plans procedure.

    - Filing- NCMRs are filed by NCMR number in a controlled location by QA personnel for easy reference.

- Reference

QA-004, Deviations

QA-018, Corrective and Preventive Action Plans

QA-026, Non-Conforming Material Process Flowchart

QAF-005, Non-Conforming Material Record (NCMR)

- Attachments

N/A

Document Revision History

| Change Order Number: | Version Number: | Description of Change: | Date: | Updated by: |
| --- | --- | --- | --- | --- |
| CC-100526-01 | 000 | New Document | 05/28/10 | A.M.M. |