- Purpose

To establish procedure for the processing of capacitive barriers including labeling, inspection, release, and inventory management.

- Scope

This procedure applies to all items relating to the processing of the capacitive barriers.

- Definitions

| Capacitive Barrier | The capacitive barrier is a plastic filter placed directly on the glass electrode surface of the ClearView device and creates a barrier between the patient’s finger and the glass electrode. |
| --- | --- |
| Capacitive Barrier Container | The container is an 8 oz. flat tin container (4” diameter x 1 ¼” high) that holds 500 pieces +/- 5% per container and includes one piece of pre-cut packing foam. |
| Packing Foam | The packing foam is a round piece of compressed foam 3 ½” in diameter and ½” thick placed between the cover lid and the top of the capacitive barriers to hold the capacitive barriers in place during shipment. |
| Labels | Labels applied to the capacitive barrier container include: - Capacitive Barrier Container Label – 3.5” in diameter, printed in color on white Avery Label 5353. - Capacitive Barrier Part Label – 1” x 2-5/8” size, printed black ink on white Avery Label 5160. Label clearly states the product specification number, revision level, quantity, and lot number. - Investigational Device Label – 1.75” x 3.25” size, printed black ink on white Avery Label 8463. |

- Responsibilities

| EPIC Employee | - Complies with the policy and procedure. - Ensures the most current version of this document is used when referenced. |
| --- | --- |
| Departmental Management | - Ensures departmental personnel are properly trained before using this policy or procedure. - Provides a timely inspection and approval/rejection of materials. |
| Executive Management | - Ensures the requirements of this plan are understood and implemented throughout all applicable levels of the company. |
| Quality Assurance | - Inspects and releases capacitive barriers. - Complies and enforces the standards within this document. - Monitors the implementation and effectiveness of this document. - Audits to ensure compliance with the referenced procedures. |

- Policy

EPIC establishes and maintains procedures to ensure that labeling, inspection, release, and inventory management of the capacitive barriers is performed to the appropriate levels as defined in controlled documentation to meet the requirements of the quality system and product specifications.

- Materials

The materials required for the processing of the capacitive barriers include:

- Avery Label 5353: Full-sheet shipping label, white, self-adhesive, 8 ½” x 11”, 100 per box (Capacitive Barrier Container Label)

- Avery Label 5160: White address label, self-adhesive, 1” x 2-5/8”, 3000 per box (Capacitive Barrier Part Label)

- Investigational Device Label, 1282-1600

- C-Line 65001: Cleer Adheer laminating film, 2.0 mil thickness/gauge, 9” x 12”, 50 per box (Capacitive Barrier Container Label)

- Color Printer

- Cutting Implement

- Procedure

7.1 Convert Stock to Inventory

    - Materials are received and handled according to the Receiving Process (OPS-004). The Receiving Clerk documents receipt of the capacitive barriers on the Receiving Log (OPSL-007), labels the materials and forwards the material to QA for quarantine until completion of the inspection process.

    - The QC Inspector processed quarantined materials according to the Inspection and Release of Incoming, In-process and Finished Devices and Other Materials (QA-016) and the Capacitive Barrier specifications (1282-1025). Product received that meets requirements is released to stock for further processing.

    - Two different labels are generated for each capacitive barrier container. For products going to be used in IDE applications, a third label is applied to the bottom side of the container (1282-1600, Investigational Device Label).

      - Record the lot number and revision level of the materials being processed on the Capacitive Barrier Device History Record, (QAF-013).

      - Capacitive Barrier Container Label

        - Print the label on a white, self-adhesive full-sheet label (Avery Label 5353).  The label is 3.5” in diameter and printed in color. Multiple labels are printed on one sheet. The content of the label is defined in Attachment A - Sample Labels.  Verify that the label information is accurate.  Record the quantity of labels printed, sign and date the Capacitive Barrier DHR record in the space provided.

        - A full-sheet of laminating film (C-Line 65001) is applied to the top of the label sheet. The laminate will protect against moisture and long term wear. Individual labels are then cut from the full-sheet printed label. Record the quantity of labels that are ultimately acceptable out of the cutting process, sign and date the Capacitive Barrier DHR record in the space provided.  Any labels that do not meet visual requirements will be shredded immediately.  Record the quantity of labels shredded, sign and date the Capacitive Barrier DHR record in the space provided.  If no labels are shredded, record zero in the space provided.

        - Place the cut label on the top side of the container cover. (see Attachment A – Sample Labels)  Record the quantity of labels applied, sign and date the Capacitive Barrier DHR record in the space provided.

      - Capacitive Barrier Part Label – This label clearly states the product specification number, revision level, quantity, and lot number.

        - Print the label with black ink on a white, self-adhesive, 1” x 2-5/8” label (Avery Label 5160). Verify that the label information is accurate (i.e., part number, revision level and lot number match the Receiving Inspection Record).  Record the quantity of labels printed, sign and date the Capacitive Barrier DHR record in the space provided.

        - Remove the existing factory label on the bottom side of the container before placing the Part Label.  Place the Part Label on the bottom side of the container. Record the quantity of labels applied, sign and date the Capacitive Barrier DHR record in the space provided.

        - If an IDE Label is also being applied, the Part Label is placed above it. (see Attachment A – Sample Labels).  Obtain the IDE Label from inventory.  Apply below the Part Number Label.  Sign and date the Capacitive Barrier DHR record in the space provided.

        - If an IDE label is not applied, cross through and initial and date the IDE Label section.

      - Upon completion of the labeling of the Capacitive Barriers, all materials will be forwarded to QC for inspection and release activities.

    - Quality Control Inspection

      - Quality Control will review the labeling to ensure that all specifications are met.  Specifically, the following characteristics will be inspected:

        - All labels are placed as indicated in Attachment A.

        - All label information is accurate, matching the layout and specifications identified in Attachment A and the Receiving Inspection Record (QAF-012).

        - Label visual quality meets these criteria:

          - Print quality is clear, not blurry.

          - Font type and size match examples in Attachment A.

          - No visual defects such as smudges, dirt marks, and/or label imperfections when viewed at 18 inches with normal light and no magnification.

      - Record the quantity of containers that meet the inspection criteria for each of the characteristics listed above on the Capacitive Barrier DHR in the space provided.

      - Record the quantity of containers that do not meet the inspection criteria in the Qty Non-Conforming space provided in the Capacitive Barrier DHR.  Initiate a Non-Conforming Material Report (QAF-011) according to the requirements specified in the Non-conforming Material procedure.  Record the NCMR number in the Notes section of the DHR.

      - Sign and date the Capacitive Barrier DHR in the space provided.  Record the total quantity of containers ready for release to inventory.

      - Forward accepted materials to the appropriate inventory location.

- Reference Documents

OPS-004, Receiving Process

OPSL-007, Receiving Log

QA-016, Release of Incoming, In-Process and Finished Services and Other Materials

QA-017, Non-conforming Material

QAF-011, Non-Conforming Material Report

QAF-012, Receiving Inspection Record

QAF-013, Capacitive Barrier Device History Record (DHR)

1282-1025, Capacitive Barrier

1282-1600, Investigational Device Label

- Attachments

Attachment A – Sample Labels

Attachment A – Sample Labels

Capacitive Barrier Container Label

Capacitive Barriers

Product #: 1282-1025     Lot #: 110829-01

www.epicdiagnostics.com

Label Specifications:

Size: 3.5 inch diameter circle

Ink: 4-color process

Material: White, self-adhesive label

Font style: Calibri

Type size: 20pt for Capacitive Barriers/all else in 9pt

Product #: 1282-1025       Revision: 000

Quantity: 500               Lot #: 110826-01

Capacitive Barrier Part Label

Label Specifications:

Size: 1” high x 2-5/8” wide rectangle

Ink: Black

Material: White, self-adhesive label

Font style: Calibri

Type size: 11pt/categories in bold type

Investigational Device Label

Sample Label

Label Specifications:

Label made to specifications approved in 1282-1600.

Document Revision History

| Change Order Number: | Version Number: | Description of Change: | Date: | Updated by: |
| --- | --- | --- | --- | --- |
| CC-110829-01 | 000 | New Document | 8/31/11 | J. Ayres |