ORIGINATOR:

| Name | Job Title | Signature | Date |
| --- | --- | --- | --- |
| Lloyd Kurth | Technology Director |  |  |

APPROVALS:

| Name | Job Title | Signature | Date |
| --- | --- | --- | --- |
| Nancy Rizzo | Founder, President and CTO |  |  |
| Brian Cullinan | Software Developer |  |  |
| Dan Miller | Manufacturing Engineer |  |  |
| Jan Ayres | Operations Director |  |  |
| Andrea Miller | Quality Assurance Director |  |  |

# Purpose

This Design Control Statement (DCS) documents the design control history for ClearView software through the release of version 2.1. This document may be revised as necessary with all revisions subject to design review.

# Scope

This DCS defines design controls for the ClearView software. The ClearView software is one component of the ClearView system. Design controls for the ClearView system, and other components, are not part of this DCS.

# Reference Documents

DDP-001, Design & Development Plan ClearView System

EG-001, Design Control Procedure

EG-011, Software Validation Procedure

EG-012, Risk Management

EG-016, Software Verification Procedure

EG-021, Software Change Control Procedure

EG-022, Software Development Guidance

Regulatory Documentation:

FDA 21 CFR 820, Quality System Regulations (QSRs)

ISO 13485:2003 Medical Devices – Quality Management Systems – Requirements for Regulatory Purposes

# Product Description

The EPIC ClearView system consists of, in part, the ClearView device and the ClearView software running on a Windows PC. Refer to DDP-001 for a full description of the ClearView system and the specific design controls mandated for ClearView software development.

The ClearView software consists of Microsoft Windows based applications and services that control the ClearView device and perform the mathematical analysis of the captured images.

# Design Control History

EPIC contracted with netFusion, a contract software development organization, for the initial development of the ClearView software. Although project reviews were conducted to assess the progress of the work performed by netFusion, these reviews were not documented.  The documentation that was generated by netFusion (controlled by their quality system) which includes detailed design specifications is included in Attachment A of this document.  A table of contents for the netFusion documentation is also included in Attachment A.  EPIC then contracted with Sage Business, an independent software development consultant, to further the design development of the ClearView software based on the design specifications provided by netFusion.  Design control authority was solely EPIC’s responsibility.  Two design reviews were conducted as a part of the work executed by Sage Business, one dated 10/27/09 and a second review conducted on 10/28/09.  These two reviews were interim reviews conducted to discuss the development work being managed by Sage Business and not formal phase reviews as currently required by EPIC’s design control procedures.

In late 2009, EPIC decided to hire staff to continue the ongoing development of the ClearView software.  From that point forward, the design history file (DHF) and all design control requirements were the responsibility of EPIC staff.  A great deal of design work was required to bring the software and the DHF documentation into compliance with EPIC’s quality system.  This design work included the creation of software requirements, initial risk analysis, and software specifications.

Additionally, verification and validation activities were performed in order to verify ClearView software performance, and approved to release software for use in clinical investigations.  The following table summarizes the design documentation to date.

| Version | Date | Reason | Notes |
| --- | --- | --- | --- |
| 0.9.2.7 |  |  | Not released |
| 0.9.7.7 | 24 September 2010 | Project Release |  |
| 1.0.0 | 24 September 2010 | Project Release | * See below. |
| 0.9.7.8 | 31 March 2011 | Patch Release | Patch release (see Attachment B for a copy of the original email release) |
| 0.9.8.0 | 24 May 2011 | Project Release |  |
| 0.9.9.1 | 24 May 2011 | Project Approval |  |
|  | 08 July 2011 | Project Release |  |
| 0.9.9.2 | 31 August 2011 | Project Approval |  |
|  | 31 August 2011 | Project Release |  |
| 1.1.0 | 20 October 2011 | Project Approval | Not released |
| 1.1.1.2 | 05 March 2012 | Project Release |  |
| 1.1.1.3 | 14 March 2012 | Project Approval | Documented on Design Change Summary Form |
|  | 16 May 2012 | Project Release |  |
| 1.1.1.4 | 30 July 2012 | Project Approval |  |
|  | 21 August 2012 | Project Release |  |
| 2.0 | 18 July 2012 | Project Approval |  |
|  | 02 October 2012 | Project Release |  |
| 2.1 | 25 October 2012 | Project Approval |  |
|  | 21 November 2012 | Project Release |  |

In 2012, work was done to create design and development plans (DDPs) to guide software development to ensure that design controls were proactively planned and met prior to release of software for use in clinical investigations.  Five separate plans were reviewed and approved and are listed in the table below.

| Document Number | Document Revision | Approved Date | Document Name |
| --- | --- | --- | --- |
| DDP-002 | 000 | 05 March 2012 | Design & Development Plan (DDP) ClearView Software |
| DDP-005 | 000 | 18 July 2012 | Design & Development Plan (DDP) ClearView 2.0 |
| DDP-008 | 000 | 30 July 2012 | Design & Development Plan (DDP) EPIC ClearView 1.1.1.3.1 |
| DDP-009 | 000 | 25 October 2012 | Design & Development Plan (DDP) EPIC ClearView 2.1 |

In early 2013, EPIC’s design control procedures were rewritten to introduce the ability to control development in “phases” and to consolidate all software design control guidance into one procedure.  All ClearView software development will be performed in accordance with the plans outlined in the system design and development plan, DDP-001.  DDP-002 rev. 000, DDP-005 rev. 000, DDP-008 rev. 000, and DDP-009 rev. 000 are considered historical plans that apply only to the design work done for the version specified.

# Project Requirements

This document is intended to clarify some of the history of the development of the ClearView software.  Section 5 of this statement provides a documentation of the history of the software development and the design controls utilized.  The ClearView system DDP, DDP-001, will specify all required design control elements for this project development.  A Table of Contents will be developed and controlled by the project leader or designee to identify all records that are included as a part of the ClearView System design history file (DHF).

- Attachment A: netFusion Documentation Table of Contents

netFusion Agile Dev – Description of Available Services

Requirements Analysis and Design Proposal – July 24, 2006

Requirements Analysis and Design Proposal – November 15, 2006

Business Agreements

Master Services Agreement Copies

EPIC Camera System Quotation

Risks and Assumptions

R0- Draft Copy

R0- Final Copy

R01- Draft Copy

R01- Final Copy

R2- Draft Copy

R2- Final Copy

R03- Draft Copy

R03- Final Copy

R04- Final Copy

R05- Draft Copy

R05- Final Copy

R06- Draft Copy

R06- Final Copy

R07- Draft Copy

R07- Final Copy

R08- Draft Copy

R08- Final Copy

U1- Draft Copy

U1- Final Copy

U2- Draft Copy

U2- Final Copy

U3- Draft Copy

U3- Final Copy

- Attachment B: 0.9.7.8 Patch Deployment Memo Dated 3/31/2011

Revision History

| Revision | Revision Date | Reason for Revision/Change Request | Revised By |
| --- | --- | --- | --- |
| 000 |  | Initial Release | L. Kurth |