ORIGINATOR:

| Name | Job Title | Signature | Date |
| --- | --- | --- | --- |
| Lloyd Kurth | Technology Director |  |  |

APPROVALS:

| Name | Job Title | Signature | Date |
| --- | --- | --- | --- |
| Nancy Rizzo | Founder, CEO |  |  |
| Brian Cullinan | Software Developer |  |  |
| Dan Miller | Manufacturing Engineer |  |  |
| Jan Ayres | Operations Director |  |  |
| Andrea Miller | Quality Assurance Director |  |  |

# Purpose

This initial risk analysis identifies the potential failure modes associated with the ClearView software component that is part of the ClearView System.

# Scope

The scope of this risk analysis is limited to the evaluation of the ClearView software failure modes.  This risk analysis is being conducted as a component (i.e., ClearView software) analysis and is a part of the ClearView System risk analysis documented in DD-006, Risk Analysis Report, ClearView System.

This initial risk analysis lists the potential failure modes identified when evaluating the ClearView software requirements documents (SR-206 and SR-207). Additional risk analysis will be performed as an ongoing effort over the life of the ClearView software development. The failure modes identified in this initial risk analysis will be further evaluated using a Failure Modes and Effects Analysis (FMEA) during ongoing design-controlled development.  The FMEA will assess the risks in compliance with EPIC Risk Management procedures to ensure that proper design control mitigations are implemented as necessary.

# Reference Documents

- DD-005, Risk Management Plan

- DD-006, Risk Analysis Report, ClearView System

- DD-010, Software FMEA

- DDP-010, Design and Development Plan ClearView Software

- EG-012, Risk Management

- EGF-009, Failure Modes and Effects Analysis (FMEA)

- SR-206, ClearView Software Requirements

- SR-207, ClearView Software Requirements, Algorithm

# Risk Assessment Process

The ClearView software development is being driven by a separate design and development plan, DDP-010.  This plan identified the need for risk analysis of the software during development.  Additionally, DD-005, Risk Management Plan, ClearView System, section 6.1, requires a FMEA to be completed during the product design.  The following risk analysis is the first step in completing a comprehensive FMEA for the ClearView software.

EG-012, Risk Management, sections 6.3.1 and 6.3.2, identifies the first step of the risk analysis as being the identification of potential hazards associated with using the system.  These hazards were assessed and documented in DD-006, Section 4. The next step in the process requires an estimation and assessment of the risk for each hazard.  To initiate that assessment, the requirements were reviewed for potential failure modes.  The following section documents the outcome of that assessment.

# Identified Failure Modes

The primary inputs to the analysis are the requirements documents, SR-206, ClearView Software Requirements and SR-207, ClearView Software Requirements, Algorithm.  After the list of potential failure modes was compiled, they were grouped into the categories in the following subsections. This grouping was done simply to organize the potential failure modes that may be related to one another. Further risk analysis during later stages of the development may re-categorize some potential failure modes.

## Use Failures

This set of failure modes would generally prevent the ClearView system from being used to produce correct scores or reports.

| Can perform operations out of order | Cannot turn image alignment display on |
| --- | --- |
| Cannot access application functions | Cannot view flattened images |
| Cannot add patient | Cannot view graphical representation of NS coefficient |
| Cannot collect new set of images for patient | Cannot view map of energized images to body locations |
| Cannot communicate with camera | Cannot view patient information |
| Cannot communicate with device firmware | Cannot view sector alignment |
| Cannot correct invalid entry | Cannot view stored scans |
| Cannot create digital copy of report | Fail to associate images with the patient |
| Cannot determine device connection status | Fail to associate scan with most recent passing setup verification |
| Cannot fully identify patient | Fail to automatically set current angle or X and Y positions |
| Cannot initiate analysis after capturing all images | Fail to detect invalid patient data prior to storage |
| Cannot modify patient information | Fail to display actual angle and midpoint values |
| Cannot modify/correct image alignment | Fail to display crosshair |
| Cannot navigate application functions | Fail to display ellipse |
| Cannot print report | Fail to display energized image after capture |
| Cannot recapture an image | Fail to display finger image used for alignment after capture |
| Cannot record extra information about patient | Fail to display image alignment on energized image |
| Cannot record whether or not patient has consumed alcohol | Fail to display image alignment on finger image |
| Cannot record whether or not patient has consumed caffeine | Fail to display instructions and status during finger image capture |
| Cannot record whether or not patient has consumed wheat | Fail to display live preview of finger image |
| Cannot resubmit images for analysis | Fail to display Y axis distinguished from X axis |
| Cannot search patients | Fail to include all organ/structure data on report |
| Cannot search reports | Fail to include date of scan on report |
| Cannot search scans | Fail to include patient biographical data on report |
| Cannot search scans by patient name | Fail to notify user of invalid patent data |
| Cannot search scans by scan date | Fail to prevent duplicate patient |
| Cannot select patient before starting a scan | Fail to prevent image analysis on incomplete image set |
| Cannot set patient birthdate for search | Fail to prompt user before analysis begins |
| Cannot set patient birthdate to add | Fail to record date of original image capture |
| Cannot set patient gender to add | Fail to retrieve stored data |
| Cannot set patient name for search | Fail to store images and alignment values |
| Cannot set patient name to add | Fail to store necessary data |
| Cannot sort scans | Fail to store scans |
| Cannot sort scans by patient name | Fail to store setup verification with correct time |
| Cannot sort scans by scan date |  |

## Setup Failures

This set of failure modes would generally prevent the ClearView system from being properly set up to produce correct scores and reports

| Cannot adjust acceptable failures | Fail to detect corona of setup verification image |
| --- | --- |
| Cannot adjust average intensity standard deviations per region | Fail to detect failed setup verification image using acceptance criteria |
| Cannot adjust bright pixels per region | Fail to detect failed setup verification image with unacceptable noise level |
| Cannot choose predefined device control set | Fail to detect failing setup verification images |
| Cannot configure brightness | Fail to detect high plateau of setup verification image |
| Cannot configure camera to connect to | Fail to detect inner section of setup verification image |
| Cannot configure exposure delay | Fail to detect number of pixels outside the acceptance criteria |
| Cannot configure exposure duration | Fail to detect outer section of setup verification image |
| Cannot configure gain | Fail to detect passing setup verification images |
| Cannot configure percent of power | Fail to detect setup verification that should fail |
| Cannot configure predefined device controls sets | Fail to detect setup verification that should pass |
| Cannot start setup verification | Fail to detect that setup verification needs to occur |
| Fail to automatically capture setup verification images | Fail to evaluate setup verification image |
| Fail to capture warm-up images | Fail to generate correct sum of differences |
| Fail to detect amount of noise in setup verification image section | Fail to pass setup verification image using acceptance criteria |
| Fail to detect area of corona in setup verification image section | Fail to pass setup verification image with acceptable noise level |
| Fail to detect brightness of setup verification image section | Fail to properly divide setup verification image into 4 sections |
| Fail to detect characteristics of setup verification image section | Fail to store setup verification |

## Calculation Failures

This set of failure modes would generally prevent the scoring algorithms from operating properly to produce correct scores and reports.

| Fail to process image in correct format | Fail to properly determine sector area |
| --- | --- |
| Fail to process mathematical functions | Fail to properly determine the ring intensity of the corona for sector |
| Fail to properly apply rules to determine final score | Fail to properly determine third form value for sector |
| Fail to properly determine average intensity of a concentric ring for sector | Fail to properly determine total score |
| Fail to properly determine average intensity of sector | Fail to properly determine z-score |
| Fail to properly determine base score for each coefficient for sector | Fail to properly divide energized finger image into sectors |
| Fail to properly determine break coefficient for sector | Fail to properly measure abrupt breaks |
| Fail to properly determine form value for sector | Fail to properly measure amount of change in pixel intensity along the degree line |
| Fail to properly determine fractal value | Fail to properly measure attributes for each concentric ring inside each sector |
| Fail to properly determine logistic regression for an organ system | Fail to properly measure entropy of sector |
| Fail to properly determine NS value | Fail to properly measure gaps around the edges of the energized ellipses |
| Fail to properly determine odds ratio | Fail to properly measure the ring thickness of the corona for sector |
| Fail to properly determine reported score | Fail to properly normalize sector area |
| Fail to properly determine second form value for sector |  |

## Security Failures

This set of failure modes would generally prevent proper implementation of security procedures to protect the ClearView system from unauthorized use.

| Cannot add user | Cannot determine who changed data |
| --- | --- |
| Cannot assign permissions to user role | Cannot uniquely identify user |
| Cannot assign privileged access rights | Fail to prevent brute-force hacking attack |
| Cannot assign user role | Fail to prevent unauthorized access |
| Cannot define user role | Fail to prevent user from accessing privileged functions |
| Cannot determine what data changed | Fail to store audit trail entries |
| Cannot determine when data changed |  |

## Ease of Use Failures

This set of failure modes would not generally prevent the ClearView system from being used to produce correct scores and reports but would hinder the ability of the user to effectively use the system.

| Cannot access reports from patient search | Cannot update user account |
| --- | --- |
| Cannot access reports from scan search | Cannot view previous successful setup verification |
| Cannot cancel function in progress | Fail to default sorting of scans from newest to oldest |
| Cannot search reports by date of scan | Fail to display captured setup verification images |
| Cannot set email address | Fail to display live preview of setup verification image |
| Cannot set patient gender for search | Fail to display patient name during image capture |
| Cannot set user's name | Fail to notify user that questions are not completed |
| Cannot switch users easily | Fail to show indication of setup verification passage/failure |
| Cannot turn image alignment display off | Fail to show setup verification progress |

## Test Facility Failures

ClearView contains a number of facilities that aid in testing device and algorithm revisions. Failure modes associated with these facilities hinder EPIC’s ability to further develop the ClearView system.

| Cannot capture a single image | Cannot search scans by indication of resubmitted data |
| --- | --- |
| Cannot clear detailed device information | Cannot search scans by software version |
| Cannot export all data | Cannot select location on disk to export data |
| Cannot export coefficient data | Cannot sort scans by indication of resubmitted data |
| Cannot export de-identified patient data | Cannot sort scans by software version |
| Cannot export images before analysis begins | Cannot view all full-sized images |
| Cannot export reported data | Cannot view data used to generate report |
| Cannot export scan images | Cannot view detailed device information |
| Cannot export setup verification data | Cannot view failed setup verification |
| Cannot export setup verification images | Cannot view raw data |
| Cannot load images manually | Cannot view sector divisions |
| Cannot load setup verification images manually | Fail to denote reports and data from resubmitted analysis |
| Cannot save a single image | Fail to export failed setup verification |
| Cannot save detailed device information | Fail to export setup verification |
| Cannot search reports by patient biographical data | Fail to include software version on report |
| Cannot search reports by software version | Fail to record software version used for analysis |

## Business Failures

This set of failure modes would generally hinder EPIC’s ability to protect its intellectual property and implement its model for revenue realization.

| Fail to customize report based on license | Fail to limit access to functionality |
| --- | --- |
| Fail to customize report based on permission | Fail to prevent access to software code |
| Fail to determine active license |  |

## Troubleshooting Facility Failures

This set of failure modes would generally hinder the ability of EPIC staff to troubleshoot and correct issues with the ClearView system.

| Cannot determine device events and commands | Fail to produce debug logs |
| --- | --- |
| Fail to determine software version |  |

# Subsequent Risk Assessment

A summarized FMEA was previously developed and documented in DD-010, Software FMEA.  This FMEA is a very high level summary of software failure modes and doesn’t assess many of the failure modes documented in this analysis.

The next step in the FMEA process will require an assessment of the level of risk for each of the identified failure modes in this document and a review of any additional failure modes documented in DD-010.  The Software FMEA document will be updated to reflect the entire software FMEA and all identified failure modes.  The appropriate Potential Effects of Failure, Potential Cause of Failure, and the Current Controls will be documented for each failure mode.  Severity, occurrence, and detection ratings will be assessed and the resulting Risk Priority Number (RPN) calculated.  Actions will be taken to ensure that all risk is minimized as required in the Risk Management procedure.  This FMEA will be developed in conjunction with the hazards documented in DD-006 and updates to the system risk analysis report and summary made as necessary.

The intended use for the ClearView System is still being developed.  As such, the software FMEA may need to be re-evaluated once the intended use is finalized to ensure that no additional risks are created as a part of the overall intended use of the software.

Revision History

| Revision | Revision Date | Reason for Revision/Change Request | Revised By |
| --- | --- | --- | --- |
| 000 |  | Initial Release | L. Kurth |