| ##### Project: | ClearView System |  |  | Project #: |  | 100 | Design Stage: |  | Initial |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| ## Design Approval Participants |  |  |  |  |  |  |  |  |  |  |
| Dept. |  |  | Signature |  | Date | ## Dept. |  | ## Signature |  | ## Date |
| Executive/CSO |  |  |  |  |  | ## Project Management |  |  |  |  |
| I.T. |  |  |  |  |  | ## Quality Assurance |  |  |  |  |
| Engineering (PADT) |  |  |  |  |  | ## Mfg. Engineering |  |  |  |  |
| ## Design Approved? ⁭Yes ⁭No Action Items are documented below. |  |  |  |  |  |  |  |  |  |  |
| ## Objectives: |  | ## Design approval should include attached documentation to fully describe the project and/or design control under review. The design objective should be written in a broad and complete manner to fully characterize the project and/or design control under review. During future stages, the complete design control documentation/requirements will continue to evolve and this approval is based solely on the stage and information available at the time of review. |  |  |  |  |  |  |  |  |
| ## This Initial Product Design Review is being conducted to address the concerns of design control compliance raised in CPAP 1017. The objective of the Initial Product Design Review is to provide approval to begin the design stage of the ClearView System in accordance with EPIC’s Design Control Procedure (EG-001). The design has been an ongoing development and this Design Review is intended to bring the Design History File into compliance. |  |  |  |  |  |  |  |  |  |  |
| Documentation of Design Information: Include major activities, assignments, and responsibilities where appropriate. List all attached design documentation under review. |  |  |  |  |  |  |  |  |  |  |
| DDP-001.000 Design & Development Plan (DDP) ClearView System DD-005.001 Risk Management Plan ClearView System DD-032.000 Product Requirements ClearView System |  |  |  |  |  |  |  |  |  |  |

ACTION ITEMS

| Description | Responsibility | Due Date |
| --- | --- | --- |
| Update the Risk Analysis Document (RAD), including DD-006 Risk Analysis Report, DD-009 System FMEA, DD-010 Software FMEA, DD-011 Process FMEA, and any applicable risk documents to be compliant with EG-001 Design Control Procedure and EG-012 Risk Management. | Dan Miller | Prior to the Design Phase Review |
| Update Project Task List to include action items from this design review. | Jan Ayres | 3/29/13 |
| Forward an updated Project Task List to Don Pegg for reference. | Dan Miller | 3/26/13 |
| Add to task list to ensure that risk assessment is in compliance with current ISO 14971 methods regarding the assessment of detection only during process risk analysis. | Andrea Miller | Prior to the Design Phase Review |