Approval Signatures:

Author:

| Department | Name | Signature | Date |
| --- | --- | --- | --- |
| IS | Andrew Mason |  |  |

Approval:

| Engineering | Lloyd Kurth |  |  |
| --- | --- | --- | --- |
| OPS | Mike Stowell |  |  |
| Exec | Nancy Rizzo |  |  |
| QA | Scott Pletzer |  |  |

Purpose

The purpose of this document is to describe revision history of the ClearView product from the initial placement into the Source Code Control System (SCCS), to the current version of the software.

- ClearView Revision History

The revision history contained in this document is an overview. For any of the versions listed in the table below, and for code submissions to which no new version was created (iterative development changes), the exact files changed as well as the exact changes can be viewed using the SCCS tool. The versions are listed from oldest to newest in the table below:

| Version # | Date Established | Changes Made to Codebase |
| --- | --- | --- |
|  | 04/15/2010 | Initial check-in of the codebase. Prior to this, the source was not controlled in any type of sccs mechanism. |
| 0.9.4.0 | 4/30/2010 | Modified the software to be able to bring up any patients images on any computer. Implemented a multi-threaded routine to produce the biofield images, so the process is much faster. Before the system would error if the scan was not performed on the machine that the report was being viewed on. Corrected the display of the bi ofields. Modified the capture window to properly display the appropriate buttons based on what role the user was. |
| 0.9.4.1 | 5/17/2010 | Completed the Biofield display page. The biofields are now shown along with all the relevant data. The user can magnify the biofields and display both the color enhanced and the original images. The system can now also pull images down for any patient regardless of where the scan was performed (previously this had a concurrency exception). |
| 0.9.5.0 | 6/7/2010 | Converted to the new Matlab routines. These new routines properly calculate the NS value used in the scoring algorithm. |
| 0.9.5.1 | 6/7/2010 | Modified exposure delay from 350 to 110, this results in much stronger and more intense images. Increased the default values for Gain and Brightness Added sector numbers when viewing the images. |
| 0.9.5.2 | 6/10/2010 | Modified the code so that the Brightness, Gain and exposure delay values could be modified through the interface. Added another tool on the menu to access a new page called "Camera Settings". Hooked in the calibration calculation to be able to display the actual calibration information on the calibration administration page. Removed a significant amount of unused “dead” code. Removed executable and generated objects from source control, this was a waste since they can be regenerated, this also made comparisons between versions more complex than it needed to be. |
| 0.9.5.3 | 6/18/2010 | Finished implementation of auto noise detection and added it throughout the application. The system now performs an analysis on the image to determine the proper level of noise to ignore during calculations. Modified code to address what appears to be a random bug with the GDI framework. The following line of code: mFrameGrabber.GrabFrame().Save(stream, System.Drawing.Imaging.ImageFormat.Bmp); Was converted into var tmImg = new Bitmap(mFrameGrabber.GrabFrame()); tmImg.Save(stream, System.Drawing.Imaging.ImageFormat.Bmp); This seems to give the GDI framework time to free the resource and not generate a really helpful "Generic Error" error message. |
| 0.9.5.4 | 6/21/2010 | Added ability to toggle crosshairs during the scanning process. Added the ability to export report data Verified ability to regenerate Matlab Dll from .m codebase. |
| 0.9.6.0 | 7/6/2010 | Modified the number of significant digits that are used for some coefficients from 1 to 2. Modified the way the standard deviation is calculated, it is now 2 pass. Added a routine to clean up files created on the filesystem when the user logs into the app. This runs as a background thread, did it at login to guarantee that files were not in use. Reordered report sections and section details Implemented new orientation process, this new process now uses a regular image of the finger in conjunction with the energized image to determine alignment. Added the related organ system logic & text to the patient report. Modified the image loading process to only allow images that were appropriately named to be selected. If the user needs 4L then only images starting with 4L will be selectable. Implemented the new calibration routine, this routine now compares all 5 rings to the new specifications supplied by Nancy. |
| 0.9.6.2 | 7/19/2010 | Corrected a problem where the size of the angle was being miscalculated. Changed routine that accepts the Matlab calculations to truncate rather than round to 4 decimal places the values. Modified the severity index calculations to take into consideration where the coefficient ranks Added some debug statements to the calculation process Added a score validator spreadsheet to the validation tools for the system. |
| 0.9.6.3 | 7/23/2010 | Modified scoring to use a Z score in the calculation. This simplified the calculation tremendously. Reverted the standard deviation to be the way it originally was. |
| 0.9.6.5 | 7/27/2010 | Modified display of NS graph ** Note, accidently skipped a version number. |
| 0.9.6.6 | 7/27/2010 | Corrected a bug where a value of NaN was being brought in from the Matlab calculations, this caused means and standard deviations to be off. This was a rare case that would occur when an image had no pixels lit in a given sector. The case is now properly handled. |
| 0.9.6.7 | 7/28/2010 | Corrected a conversion error that was resulting in a 1/1000 point differential in a calculation on occasion. |
| 0.9.7.0 | 9/13/2010 | Completed biofield mapping (initial design). Modified scoring routine to implement latest coefficient. |
| 0.9.7.1 | 9/15/2010 | Modified calibration routine. Added calibration fail-safe switch |
| 0.9.7.3 | 9/16/2010 | Removed unnecessary biofield total value display |
| 0.9.7.4 | 9/17/2010 | Modified calibration tolerances Added new EPIC logo to application |
| 0.9.7.5 | 9/17/2010 | Corrected bug on admin page where the detail image dialog would not display. |
| 0.9.7.6 | 9/17/2010 | Replaced organ system text in analysis report |
| 0.9.7.7 | 9/20/2010 | Corrected some misspellings in the biofields |
| 0.9.7.8 | 10/12/2010 | Added the image alignment circle to the finger image as well as the energized images in order to help the user better discern the angle. This change will not affect the calculations or values in any way; it is strictly a visual change. Added the time to the visit date that appears on the ClearView report. |
| 0.9.8.0 | 4/5/2011 | Made changes to the Camera testing utility based on feedback from the Firmware engineer. Added new Gather Capture process Modified the calibration manager to be able to just regenerate the Matlab files and bypass the comparison process (Admin use only) Modified the view capture dialog to more efficiently do resubmissions. Corrected application timeout algorithm to properly function, and to display the status to the user at the bottom of the application. Completed a working resubmission process. Committing to Perforce in order to make continuing development safer. |
| 0.9.9.0 | 5/4/2011 | New calibration process Finger image persistence Quick export facility Some overall code reorganization (more modularity) Removal of old export functionality from admin pages Removal of image analysis dialog from admin pages Added connectivity indicators to the footer of the app Added treatment IDs to the search capture dialog |
| 0.9.9.1 | 7/6/2011 | Added new release files containing a colorization change for failed calibration images. Updated calibration routine, now uses a new more precise comparison process. Updated Matlab error handling. Corrected cache timeout issue. Added the ability to save both failed and successful calibration images to disk. This capability is only for administrators. Corrected a bug that would cause an error if a scan was viewed that had no finger images. Corrected a bug where the user options were missing from the main menu. Disabled help button per quality request. Massive code cleanup. |
| 0.9.9.2 | 08/19/2011 | Upgraded runtime from .Net 2.0 to .Net 4.0 More code cleanup Added ability to save calibration log data to be used in performance assessment. Correction in the export calibration results routine. Modified the way images are handled to be more memory efficient Corrected a display bug in the biofield area. Upgraded to Crystal Reports for reporting process. |
| 1.1.0.0 | 10/31/2011 | Added Naive Bayes Scoring model Added Logistic Regression Scoring model Added Updated EPIC Scoring model Added new licensing model that supports 3 product release levels (CV and Basic). Updated application logging routine. Updated ClearView report with a completely new look. |
| 1.1.1.0 | 12/9/2011 | Updated scoring algorithm, modified report layouts. Updated licensing model code. Updated Matlab files and error handling process. |
| 1.1.1.a* | 12/29/2011 | Modified LR scoring algorithm, to correct a minor calculation issue. Added ability to view calculation coefficients in administrative section. Added a new administrator type 'techadministrator' Relabeled 2 biofields Modified code to allow different sorting of reports when using different licenses. Added the ability to have 3 license modes rather than 2. |
| 1.1.1.b | 1/12/2012 | Corrected a display issue Corrected stored procedure for searching by User Id Corrected issue on Basic report (missing value) Fixed a UTC timestamp issue Report text changes |
| 1.1.1.c | 1/13/2012 | Corrected a spelling issue on a biofield Corrected non-conformance (3 biofields with the same score values) |
| 1.1.1.d | 1/16/2012 | Modified text on report to reflect new 0-5 scoring for the Naïve Bayes calculation. Modified Naïve Bayes score to be a value of 0-5 rather than a value of 0-25. |
| 1.1.1.e | 1/31/2012 | Corrected timestamp directory name issue in export process Consolidated age calculation process into a single function Updated SQL upgrade script |
| 1.1.1.f | 2/24/2012 | Removed most demographic fields, added auditing to the ones that were left. This resulted in the following bugs being closed in the bug tracker: EPIC-45, EPIC-67, EPIC-68, EPIC-46, EPIC-39, EPIC-37, EPIC-73, EPIC-86, EPIC-87, Added device testing page in the Settings section of the application Added login tracking to prevent a user from attempting to log in unsuccessfully more than three times in a row. Added ability to track what version of device firmware was used to capture images. Fixed bug EPIC-94 |
| 1.1.1.g | 2/24/2012 | Added an option to display the plotted values on the NS page, this is only available to administrator and techadministrators. Corrected a bug where the administrator view of the ‘Heart (muscle)’ biofield was not displaying. Changed the version number in the executable from 1.1.1.f to 1.1.1.g. Added EPIC.Auditing source file to SCCS. |
| 1.1.1.h | 3/1/2012 | Corrected EPIC-97 missing insertion date in Treatment table (for audit use). Corrected save issue on settings page. Incremented the version number to 1.1.1.h Corrected EPIC-96 (NS Graph Bug) |
| 1.1.1.2 | 3/12/2012 | Stamping of release value into executables (1.1.1.2). |
| 1.1.1.2.a | 4/13/2012 | Updated timing method for exposure capture to use a timer rather than sleeping. This was done in the calibration process as well as the capture process. Fixed a bug in the Camera Tester that was preventing an exposure greater than .5 seconds from being tested (EPIC-103). Added a patent pending statement to the application. Moved the license type identifier from the header of the application to the footer. Updated the text of the report per NAMSA and Nancy. |
| 1.1.1.3 | 5/15/2012 | Final stamping of version 1.1.1.3 |

*Implemented a method for tracking changes to the software code during validation testing.  The last digit in the version number (for this round, the “2” in version 1.1.1.2) is replaced with an alpha character.  This alpha character is then revised for any coding changes implemented during the verification and validation test cycle.  
Document Revision History

| Version Number: | Description of Change: | Date: | Updated by: |
| --- | --- | --- | --- |
| 000 | Introduction | 3/5/12 | A. Mason |
| 001 | Add final release of version 1.1.1.2, test version 1.1.1.2.a and final release of version 1.1.1.3. | 5/16/12 | A. Mason |