ORIGINATOR:

| Name | Job Title | Signature | Date |
| --- | --- | --- | --- |
| Lloyd Kurth | Technology Director |  |  |

APPROVALS:

| Name | Job Title | Signature | Date |
| --- | --- | --- | --- |
| Nancy Rizzo | Founder, CEO |  |  |
| Brian Cullinan | Software Developer |  |  |
| Dan Miller | Manufacturing Engineer |  |  |
| Jan Ayres | Operations Director |  |  |
| Andrea Miller | Quality Assurance Director |  |  |

# Purpose

This DDP outlines the plan for developing, testing, and releasing the ClearView software component of the ClearView system for commercial or investigational device exemption (IDE) use by EPIC Research & Diagnostics, Inc. (EPIC). The software component interfaces with the ClearView device to control the functions of the device. The software component also contains all the algorithms for data analysis and the database for storing all data.

This document is subordinate to the ClearView System DDP (DDP-001).  DDP-001 is primarily focused on the development of the ClearView device. However, it is the overarching plan for the entire system and this plan fits within it and provides the detailed plan and control for the software.

# Scope

This DDP outlines the design control activities required for the systematic development and assessment of the ClearView software component. Systematic development and assessment requires design inputs and outputs, software design definition, implementation (writing software code), verification, validation, change control, and release, along with the appropriate reviews and oversight.

This project will document the implementation of the product requirements into software. Included in the implementation is the realization of the proprietary algorithms for analyzing scanned images to produce an assessment of the patient’s health. Initially, the algorithms will be inherited from a previous version of the software. Subsequently, algorithm modifications are part of this development plan and will be managed using change control as other software modifications are managed.

This DDP is supported by DD-033, Design Control Statement, ClearView Software, which documents the design control history for ClearView software to the present time and summarizes the design controls going forward through commercialization of the ClearView system.

# Reference Documents

This DDP makes specific reference to the following documents:

DD-005, Risk Management Plan

DD-013, ClearView Product Requirements - Software

DD-032, Product Requirements, ClearView System

DD-033, Design Control Statement, ClearView Software

DDP-001, Design & Development Plan, ClearView System

EG-001, Design Control Procedure

EG-011, Software Validation Procedure

EG-012, Risk Management

EG-016, Software Verification Procedure

EG-021, Software Change Control Procedure

EG-022, Software Development Guidance

SR-206, ClearView Software Requirements

SR-207, ClearView Software Requirements, Algorithm

Regulatory Documentation:

FDA 21 CFR 820, Quality System Regulations (QSRs)

ISO 13485:2003 Medical Devices – Quality Management Systems – Requirements for Regulatory Purposes

# Product Description

The EPIC ClearView system consists of the ClearView device and the ClearView software running on a Windows PC. Refer to DDP-001 for a full description of the ClearView system.

ClearView software consists of Microsoft Windows based applications and services that include the following functionality:

Manages user accounts; users are assigned to roles and granted permissions to perform particular functions within ClearView.

Manages records for patients, including demographic data required for the scoring algorithms.

Controls the device, sending commands to the onboard firmware via a USB interface.

Calibrates the device; calibration data provides a baseline for scoring algorithms.

Captures images via a USB interface to the onboard camera.

Captured images are analyzed using proprietary algorithms to create a report that includes summary body system scores and detailed data for specific body organs.

Stores all collected images, raw data, scores, and patient data in a database.

# Project Requirements

## The ClearView system is expected to be classified as an FDA Class II medical device and the ClearView software will be required to comply with all elements of EG-001, Design Control Process.

## This DDP provides a high-level plan for design and development of the ClearView software. Detailed plans are maintained in a separate Project Task List managed by the Operations Director. Process Review Meetings will be conducted monthly to review and update project activities.

## Product requirements for software are documented in SR-206, ClearView Software Requirements and SR-207, ClearView Software Requirements, Algorithm. These requirements are the inputs that drive the design of the ClearView software. Verification and validation plans and reports will be reviewed to demonstrate that the product requirements have been met.

## Design outputs will be traced back to product requirements at appropriate stages of development and fully documented in a Traceability Analysis Document.  A current version of the traceability analysis is documented in DD-013, ClearView Product Requirements- Software.

## Verification of the ClearView software will be performed according to EG-016, Software Verification Procedure. (See Section 9 for more details.)

## Validation of the ClearView software will be performed according to EG-011, Software Validation Procedure. (See Section 9 for more details.)

## Risk management is controlled by DDP-001 in compliance with EG-012, Risk Management.   Risk analysis performed during software development will follow the plan outlined in DD-005, Risk Management Plan.  (See Section 10 for more details.)

## Software changes will be managed according to EG-021, Software Change Control Procedure.

## Each released version of the ClearView software will be managed as a phase, as defined by EG-001. Software verification activities that can be performed without a ClearView device will be controlled by this DDP. When those verification activities are complete, reviewed, and approved, remaining verification activities that require a ClearView device will be performed under control of DDP-001. All versions will be released for validation and/or use in clinical investigations under DDP-001.

## A record of every software version submitted for verification and validation will be maintained in the Revision History Record. Entries in the record for versions submitted for verification will be approved at a Phase Design Review. Entries for versions submitted for validation will be approved at a Phase Release Review.

# Responsibilities

## Information Technology (IT)

Design, development, verification, and release of ClearView software

Software risk analysis

Installation of software for verification and validation

## Chief Science Officer (CSO)

Design, verification, and validation of proprietary scoring algorithms for analyzing images

Software risk analysis

## Quality Assurance (QA)

Approval and oversight of design control activities for ClearView software

Software risk analysis

Development and execution of software validation plans and protocols

## Operations

Maintaining Project Task List and monitoring progress of project and business goals

## Manufacturing

Approval and review of initial design, and design changes, related to how the software interfaces with the ClearView device

## ClearView System Program Manager

Oversight of design and development to ensure compliance with EPIC’s quality system and to all applicable standards and regulations

## Software Project Leader

Day-to-day management and oversight of the ClearView software development according to this DDP

Reporting of project status to team members, Project Implementation, and EPIC management

Software risk analysis

## Hardware Project Leader

Oversight of the ClearView system development according to DDP-001

Software risk analysis

# Organization, Interfaces, and Approval Authority

## The ClearView software is one component of the ClearView system. The ClearView system plan is DDP-001 and this DDP is included by reference. The ClearView device is also a component of the ClearView system and its development is covered by DDP-001.

## Product versions that include software and device changes that are interdependent require coordination by the Software Project Leader and Hardware Project Leader. The ClearView System Program Manager has the overall responsibility for version release.

## Approval authorities are defined by the following matrix:

| Name | Title/Role | Company | Review Approval Required |
| --- | --- | --- | --- |
| Nancy Rizzo | Founder and CEO, CSO ClearView System Program Manager | EPIC | Design Review Board Member Software development documentation approvals |
| Lloyd Kurth | Technology Director Software Project Leader | EPIC | Design Review Board Member Software Design Evaluation Software Verification Review Software development documentation approvals |
| Jan Ayres | Operations Director | EPIC | Design Review Board Member Project Task List management |
| Andrea Miller | Quality Assurance Director | EPIC | Design Review Board Member Software Design Evaluation Software Verification Review Software development documentation approvals |
| Dan Miller | Manufacturing Engineer Hardware Project Leader | EPIC | Design Review Board Member Software Verification Review |
| Brian Cullinan | Software Developer | EPIC | Design Review Board Member Software Design Evaluation Software Verification Review Software development documentation approvals |

# Major Tasks & Milestones

Software Requirements

Initial Risk Analysis

Product Design Review

Multiple phases, managed by change control, each resulting in a software version, each consisting of:

Investigation, prototypes, planning

Create/update verification and validation plan

Software Design

Software Design Evaluation

Software Construction, including unit tests

Code Review, including unit tests

Review and update of software risk analysis (i.e., FMEA)

Phase Design Review

Software Verification (see note below)

Review and update of software risk analysis (i.e., FMEA)

Software Verification Review (see note below)

System Verification (controlled by DDP-001)

Phase Release Review (controlled by DDP-001)

(Remaining tasks and milestones controlled by DDP-001)

NOTE: Software verification activities that can be performed without the use of a ClearView device may be performed and reviewed prior to System Verification. Following these activities, a design review will be conducted to evaluate the verification results and to ensure that all portions of the DDP are satisfied. Following a successful review, System Verification will be controlled by DDP-001.

# Design Verification & Validation

A separate verification and validation plan will be established to plan the specific activities required for software verification and validation.

## Verification

Software verification consists of the verification activities as defined by EG-016, Software Verification Procedure. Most software verification activities can be planned and controlled under this DDP. Some software verification tests require the use of a ClearView device. These tests will be part of the overall verification plan, but the execution of these tests will be controlled by DDP-001.

## Validation

Validation of the ClearView software will be performed according to EG-011, Software Validation Procedure. Validation requires use of the ClearView device and will be controlled by DDP-001.

## A verification and validation plan will be developed prior to any verification or validation activities.  Since the validation will be controlled by DDP-001, the verification and validation plan will require review by those individuals identified in DDP-001 and in section 7.3 of this DDP.

# Risk Management

The risk management plan for ClearView system is defined by DD-005, Risk Management Plan. Software risk management activities are part of that plan. The following activities specifically related to software will be controlled under this DDP.

## An initial risk analysis for software will be performed prior to the Product Design Review. A Failure Mode and Effects Analysis (FMEA) will be initiated to identify product functions and processes and associated potential failure modes.

## Product functions and processes and potential failure modes will be added and revised as necessary during development as new discoveries are made and new information is available.

## Severity, occurrence rating, and likelihood of detection values will be assigned to each potential failure mode in accordance with EPIC’s Risk Management procedure. These values will be reassessed when current controls change during development.

## The FMEA will be reviewed and reassessed during Software Design and Software Construction to ensure identified risks are addressed.

## The FMEA will reference specific verification and validation tests and activities to identify risk mitigation steps. Verification and validation reports will ensure mitigation steps have been completed.

## The FMEA will be reviewed and any updates approved at each Phase Release Review. The Design Review Board is responsible for reviewing and approving the FMEA.

# Project constraints

Known project constraints at this time are identified and maintained in the Project Task List.  Any additional constraints will be managed through EPIC staff and process review meetings and/or documented via revisions of this plan.

# Design history file (DHF)

The outputs of this effort will be generated and added to the Design History File (DHF) in accordance with the requirements of EG-001, Design Control Procedure.

Revision History

| Revision | Revision Date | Reason for Revision/Change Request | Revised By |
| --- | --- | --- | --- |
| 000 |  | Initial Release | L. Kurth |