ORIGINATOR:

| Name | Job Title | Signature | Date |
| --- | --- | --- | --- |
| Brian Cullinan | Software Developer |  |  |

APPROVALS:

| Name | Job Title | Signature | Date |
| --- | --- | --- | --- |
| Nancy Rizzo | Founder, CEO |  |  |
| Lloyd Kurth | Technology Director |  |  |
| Dan Miller | Manufacturing Engineer |  |  |
| Jan Ayres | Project Implementation Manager |  |  |
| Andrea Miller | Quality Assurance Director |  |  |

Revision History

| Revision | Revision Date | Reason for Revision/Change Request | Revised By |
| --- | --- | --- | --- |
| 000 | 18-Sep-12 | Initial Release | L. Kurth B. Cullinan |

Table of Contents

1 Purpose	3

2 Scope	3

3 Definitions	3

3.1 References	3

3.2 Terms	3

4 Summary	3

5 Results	4

6 Conclusions	4

7 Attachments	4

# Purpose

The purpose of this document is to describe the steps that were taken to verify the setup verification (calibration) process created by Geost performs properly.

# Scope

The scope of this document only covers the verification activities related to calibration.  This includes all calibration testing necessary for ClearView and the underlying mathematics taking place when a calibration is performed.

This document does not cover any verification activities performed for other applications functions.  This does not cover the function between storing verification data in the database.  This does not cover how verification images are used within the finger analysis coefficient algorithm.

# Definitions

# Summary

When EPIC entered development of ClearView 2.0, it was determined that the current calibration process could be improved to provide a tighter, more testable analysis of calibration images.  This would allow EPIC to enter hardware activities such as sourcing a new camera and glass electrode and maintain a defined control over the output of changing hardware.

The previous calibration routine helps to identify camera brightness, gain, and exposure settings that produce images with expected analysis results and filter out the settings with unexpected results.  It did this by grouping pixels with a certain brightness level, finding the sum of all pixels within that range, and determining if the sum of those pixels falls within the specified range of acceptable pixels with that brightness.

The new Geost algorithm takes an input calibration image, compares the pixels to a baseline image of many calibration images averaged together.  The number of pixels that fall within the standard deviate from the average image are summed together and counted as passing pixels.  The number of pixels that fall outside of the standard deviation are considered failing.  The sum of each ring is considered to determine if the image passing or failing.  All rings must pass, as well as the total count of passing pixels, in order for the image to be considered passing.

It is arguable that the two functions do exactly the same thing, which is arriving at the same conclusion, using two different methods.  The tests detailed below will further prove that.  Tests will also be performed to show that given a set of new passing images does not adversely affect the outcome of the patient results.  This will be done by creating new analysis records based on specific known patient results, swapping out calibration images, and viewing the new analysis results.

# Results

During the load of known passing and failing images, 10 images were found marked as failing that passed calibration on 1.1.1.2.  These images will be included in the set of passing images, making the result set 90 failing and 90 passing.

Out of all images, none were found to fail calibration in 1.1.1.2 and passing calibration in 2.1.  That is, all images marked as failing also failed 2.1.

33 images marked as passing by the routine in 1.1.1.2 failed in 2.1 due to a tighter restriction defined around the High Plateau region of the image.

1 image failed due to restrictions places on the inner circle on the image (intended to be black).  This is due to some amount of noise or light in the image.

# Conclusions

The results show that all the tests have the expected outcome and the new calibration routine performs just as well as the previous routine.

# Attachments

Geost Setup Verification Procedure

List of passing and failing images