ORIGINATOR:

| Name | Job Title | Signature | Date |
| --- | --- | --- | --- |
| Brian Cullinan | Software Developer |  |  |

APPROVALS:

| Name | Job Title | Signature | Date |
| --- | --- | --- | --- |
| Nancy Rizzo | Founder, CEO |  |  |
| Lloyd Kurth | Technology Director |  |  |
| Dan Miller | Manufacturing Engineer |  |  |
| Jan Ayres | Project Implementation Manager |  |  |
| Andrea Miller | Quality Assurance Director |  |  |

Revision History

| Revision | Revision Date | Reason for Revision/Change Request | Revised By |
| --- | --- | --- | --- |
| 000 |  | Initial Release | B. Cullinan |

# Purpose

The purpose of this procedure is to create a data set for use in the verification of changes to the setup verification analysis used in the ClearView software to verify ClearView System setup prior to use.

# Scope

This protocol provides instructions for performing verification activities for Setup Verification functions of the ClearView software.  Additional verification and validation activities may be specified in additional protocols.

# Summary

Previous version of software design documentation referred to calibration image evaluation.  The activity performed verified the device performance using a quantitative analysis of the images generated.  The analysis is not intended to correct the graduation of the measurement device or tune the device controls.  Instead, the setup verification is intended to verify that equipment setup factors like, cleaning of the glass, camera settings (i.e., brightness and gain), and system warm-up.   In order to more clearly describe the function to users and reduce confusion, the name of this activity has been changed from calibration to setup verification which is intended to convey the purpose of the activity more clearly to users.

The previous setup verification routine analysis grouped pixels with a certain brightness level, finding the sum of all pixels within that range, and determining if the sum of those pixels falls within the specified range of acceptable pixels with that brightness.  The acceptance range was based on standard deviations from expected average pixel brightness for each pixel in the 320 x 240 bitmap.

EPIC contracted with a vendor, Geost, to develop an analysis mechanism that would provide an analysis based on a compilation of acceptable calibration images (called the baseline image).  Geost provided a mechanism [add documentation reference here] which assesses the collected setup verification image to the baseline image.  The assessment compares

each pixel brightness value in the setup verification image to the coordinating baseline image pixel brightness.  Any pixel brightness that falls outside of the selected standard deviation from the average pixel brightness of the baseline image fails.  The total number of failed pixels within each of four concentric rings (starting at the inner corona and moving outward) is then used to determine whether the number of failed pixels is below a threshold limit.  Any ring that exceeds the threshold limit of failed pixels will cause the image to fail the analysis.  Additionally, the total count of failed pixels within the entire image is compared to a separate threshold limit.  If the total count of failed pixels is higher than the total pixel count threshold, the image fails the analysis.

The old and the new analysis perform a similar assessment in that they are both examining pixel brightness.  This protocol is intended to demonstrate that the new analysis does not adversely affect the outcome of the patient results and provides a similar analysis to the existing analysis.  The protocol executes utilizes existing patient image data (both setup verification images and patient images) and runs them through the new setup verification analysis.

# Procedure

The following steps will be performed in order to collect all necessary data.

## Overall Instructions

  - Follow the instructions provided within this protocol as written.   Any deviations from the written protocol will be recorded on Attachment A, Deviations from Protocol.

    - QA will oversee the execution of the protocol.  A QA designee will review each deviation as documented prior to moving forward in the test process.

    - QA will be required to determine whether the deviation requires formal documentation through EPIC’s Deviation procedure or if the deviation is minor enough to warrant documentation only.

    - Upon completion of the protocol, the report approvers have the responsibility to review all deviations recorded to determine whether or not they significantly impact the test process.  Any deviation deemed significant will be handled through EPIC’s Deviations procedure.

    - The individual performing the test protocol will complete each column of Attachment A for all deviations from the written protocol prior to moving forward.  QA will review each deviation prior to moving forward in the test process.  This review is indicated by documenting a signature and date in the Review column of the Deviations from Protocol worksheet located in Attachment A.  Additional copies of Attachment A may be used if necessary.

  - The written protocol is intended to capture the steps needed to properly execute the tests.  However, given the evolving nature of software development, the instructions may not be 100% accurate or complete.  Therefore, any minor deviation from the written instructions will be corrected in writing during the execution of the test protocol.  These corrections will be reviewed by QA as a part of the test results analysis.  The review will determine appropriate corrective and preventive action for any deviations notes.

  - Record the results of each test step by initialing and dating in the space provided.  If the result cannot be verified, record a non-conformance reference number on the protocol and the Non-Conformances Worksheet (Attachment B) and describe the failure in a specific and complete manner by completing all columns on the Non-Conformance Worksheet.

  - ClearView Error Messages

## Any error messages received will be documented and attached to the final test protocol.

## Version 1.1.1.2 Data Collection

IT will load a version of ClearView 1.1.1.2 and document the installation per the appropriate quality system procedures.  The appropriate installation documentation will be attached to this protocol.

- Attached by: 						Date:

The test designee will use the ClearView Instructions for Use (CS-100 or CS-101) to execute all instructions in this test procedure.  Log into the ClearView software v. 1.1.1.2 as an administrator.

- Verified	 Not Verified	Non-Conformance Ref. No.

Open the Calibration window and

load the first ten (10) of the known failing images from the folder: I:\Calibration_Testing\APR2011\NewCal_Images\NewCalAndyFail .  Initiate the calibration analysis.

- Verified	 Not Verified	Non-Conformance Ref. No.

Upon completion of the calibration analysis, export the calibration data when prompted and label “1112 Set 1”.

Repeat steps 3 through 5 for all 100 images (ten sets of ten images) labeling each set with the next sequential number.  Record your initials and date for each set after all data is saved locally.

| - Images |  | - Set |  | - Initial and Date |
| --- | --- | --- | --- | --- |
| - SN07_01_XX through SN07_10_XX |  | - 1112 Set 1 |  |  |
| - SN07_11_XX through SN07_20_XX |  | - 1112 Set 2 |  |  |
| - SN07_21_XX through SN07_30_XX |  | - 1112 Set 3 |  |  |
| - SN07_31_XX through SN07_40_XX |  | - 1112 Set 4 |  |  |
| - SN07_41_XX through SN07_50_XX |  | - 1112 Set 5 |  |  |
| - SN07_51_XX through SN07_60_XX |  | - 1112 Set 6 |  |  |
| - SN09_01_XX through SN09_59_XX |  | - 1112 Set 7 |  |  |
| - SN25_02_XX through SN25_15_XX | - 1112 Set 8 |  |  |  |
| - SN25_17_XX through SN25_29_XX | - 1112 Set 9 |  |  |  |
| - SN25_32_XX through SN25_58_XX | - 1112 Set 10 |  |  |  |

Locate the saved calibration data sets.  Compile all of the results in to a single spreadsheet.

Remove the extra columns so the spreadsheet displays only the failed calibration data.  An example of the spreadsheet looks like this:

| Image Filename | Failed Pixels | Status |
| --- | --- | --- |
| Image#X.bmp | 7622 | Failed |
| … | … | … |
| Image#X.bmp | 6993 | Failed |

- Print out a final version of the results, sign and date the printout and attach them to this protocol.

- Attached by: 						Date:

Repeat steps 3 through 5 for all 80 known passing images (eight sets of ten images) from the folder:

- I:\Calibration_Testing\APR2011\NewCal_Images\NewCalAndyPass.  Label each set with the next sequential number.  Record your initials and date for each set after all data is saved locally.

| - Images | - Set | - Initial and Date |
| --- | --- | --- |
| - SN09_03_XX through SN09_13_XX | - 1112 Set 11 |  |
| - SN09_14_XX through SN09_26_XX | - 1112 Set 12 |  |
| - SN09_27_XX through SN09_37_XX | - 1112 Set 13 |  |
| - SN09_38_XX through SN09_48_XX | - 1112 Set 14 |  |
| - SN09_49_XX through SN09_60_XX | - 1112 Set 15 |  |
| - SN25_01_XX through SN25_30_XX | - 1112 Set 16 |  |
| - SN25_31_XX through SN25_46_XX | - 1112 Set 17 |  |
| - SN25_47_XX through SN25_60_XX | - 1112 Set 18 |  |

Locate the saved calibration data sets.  Compile all of the results into the spreadsheet generated in step 6 above.  Remove the extra columns so the spreadsheet displays only the failed calibration data.  An example of the spreadsheet looks like this:

| Image Filename | Failed Pixels | Status |
| --- | --- | --- |
| Image#X.bmp | 7622 | Failed |
| … | … | … |
| Image#X.bmp | 6993 | Failed |

- Print out a final version of the results, sign and date the printout and attach them to this protocol.

- Attached by: 						Date:

## Version 2.1 Data Collection

- IT will load a version of ClearView 2.1 and document the installation per the appropriate quality system procedures.  The appropriate installation documentation will be attached to this protocol.

- Attached by: 						Date:

- Log into the ClearView v. 2.1 as administrator.

- Verified	 Not Verified	Non-Conformance Ref. No.

Open the Calibration window and load the first ten (10) of the known failing images from the folder: I:\Calibration_Testing\APR2011\NewCal_Images\NewCalAndyFail .  Initiate the calibration analysis.

- Verified	 Not Verified	Non-Conformance Ref. No.

Upon completion of the calibration analysis, export the calibration data when prompted and label “21 Set 1”.   Repeat steps 3 through 5 for all 100 images (ten sets of ten images) labeling each set with the next sequential number.  Record your initials and date for each set after all data is saved locally.

| - Images | - Set | - Initial and Date |
| --- | --- | --- |
| - SN07_01_XX through SN07_10_XX | - 21 Set 1 |  |
| - SN07_11_XX through SN07_20_XX | - 21 Set 2 |  |
| - SN07_21_XX through SN07_30_XX | - 21 Set 3 |  |
| - SN07_31_XX through SN07_40_XX | - 21 Set 4 |  |
| - SN07_41_XX through SN07_50_XX | - 21 Set 5 |  |
| - SN07_51_XX through SN07_60_XX | - 21 Set 6 |  |
| - SN09_01_XX through SN09_59_XX | - 21 Set 7 |  |
| - SN25_02_XX through SN25_15_XX | - 21 Set 8 |  |
| - SN25_17_XX through SN25_29_XX | - 21 Set 9 |  |
| - SN25_32_XX through SN25_58_XX | - 21 Set 10 |  |

Locate the saved calibration data sets.  Compile all of the results in to a single spreadsheet.  Remove the extra columns so the spreadsheet displays only the failed calibration data.  An example of the spreadsheet looks like this:

| Image Filename | Failed Pixels | Status |
| --- | --- | --- |
| Image#X.bmp | 7622 | Failed |
| … | … | … |
| Image#X.bmp | 6993 | Failed |

- Print out a final version of the results, sign and date the printout and attach them to this protocol.

- Attached by: 						Date:

Repeat steps 3 through 5 for all 80 known passing images (eight sets of ten images) from the folder:

- I:\Calibration_Testing\APR2011\NewCal_Images\NewCalAndyPass.  Label each set with the next sequential number.  Record your initials and date for each set after all data is saved locally.

| - Images | - Set | - Initial and Date |
| --- | --- | --- |
| - SN09_03_XX through SN09_13_XX | - 21 Set 11 |  |
| - SN09_14_XX through SN09_26_XX | - 21 Set 12 |  |
| - SN09_27_XX through SN09_37_XX | - 21 Set 13 |  |
| - SN09_38_XX through SN09_48_XX | - 21 Set 14 |  |
| - SN09_49_XX through SN09_60_XX | - 21 Set 15 |  |
| - SN25_01_XX through SN25_30_XX | - 21 Set 16 |  |
| - SN25_31_XX through SN25_46_XX | - 21 Set 17 |  |
| - SN25_47_XX through SN25_60_XX | - 21 Set 18 |  |

Locate the saved calibration data sets.  Compile all of the results into the spreadsheet generated in step 6 above.  Remove the extra columns so the spreadsheet displays only the failed calibration data.  An example of the spreadsheet looks like this:

| Image Filename | Failed Pixels | Status |
| --- | --- | --- |
| Image#X.bmp | 7622 | Failed |
| … | … | … |
| Image#X.bmp | 6993 | Failed |

- Print out a final version of the results, sign and date the printout and attach them to this protocol.

- Attached by: 						Date:

Create an Acceptance Result column.  Copy the following equation into each row of calibration data in the spreadsheet.

- =IF(AND(EXACT(Y4,"Passed"),EXACT(I4,"Passed"),EXACT(M4,"Passed"),EXACT(Q4,"Passed"),EXACT(U4,"Passed"),EXACT(Y4,"Passed"),EXACT(AC4,"Passed"),EXACT(AF4,"Passed"),EXACT(AH4,"Passed"),EXACT(AJ4,"Passed"),EXACT(AL4,"Passed")),”Passed”,”Failed”)

- Verified	 Not Verified	Non-Conformance Ref. No.

# Attachments

Attachment A, Deviations from Protocol

Attachment B, Non-conformances Worksheet

# Attachment A

Deviations from Protocol

NOTE:  This form is used for minor deviations from the protocol as written.  Fill out all sections of this form prior to moving forward in the validation process.

| No. | Date | Name | Description of Deviation (include reference to the protocol section) | Resolution/Action Taken | Reviewed |
| --- | --- | --- | --- | --- | --- |

Attachment B - Non-conformances Worksheet

NOTE:  This form is used for all protocol steps which did not perform as expected.  Fill out all sections of this form prior to moving forward in the validation process.

| No. | Date | Name | Description of Non-conformance (include reference to the protocol section) | Resolution/Action Taken |
| --- | --- | --- | --- | --- |