Originator:

| Department | Name | Signature | Date |
| --- | --- | --- | --- |
| ENG | Dan Miller |  |  |

Approval Signatures:

| Department | Name | Signature | Date |
| --- | --- | --- | --- |
| QA | Scott Pletzer |  |  |
| OPS | Mike Stowell |  |  |
| Exec | Nancy Rizzo |  |  |
| IS | Andrew Mason |  |  |

- Purpose

To determine the effect of modifying calibration acceptance limits on ClearView report scores.

- Verification Method

This study is intended to collect data to support the justification for modification of the calibration pass/fail limits.  The data is intended to show that the change in calibration values does not have a negative effect on the actual human subject values reported on the ClearView report.  The Naïve Bayes (NB) values are selected as the comparison values because they are a global check on the algorithm as a whole.

The NB scores range from 0-5 and are sensitive to the tenth decimal place. Due to the range of the scores and the knowledge of the sensitivity of the NB scores with respect to the odds ratios for each organ system a difference of 0.5 or less across the three scores is considered acceptable.  This allows us to stay within the 95% confidence limit for each organ system.

Any nonconformance or deviation from performing all sections of this validation must be resolved with appropriate written justifications prior to final approval of the validation report.

The test method employed is as follows:

Original report scores from five subjects tested in device validation procedure ENG-017.002 (device SN0986) will be compared to the scores using the same finger images but different calibration images.  The calibration images will be from the revised software using three different brightness and gain settings; one at nominal, one at the lower limit of acceptance, and one at the upper limit of acceptance.

- Procedure

  - Record system and software information in Table 1.

  - Select five subjects per section 2.0, print and label ClearView reports.

  - Calibrate device using brightness setting of 85 and gain setting of 58.  After successful calibration, save images to folder on the network labeled NewCalLimitsTesting.  Load finger images from device validation testing for first subject.  Modify angles as required and submit for analysis.  After analysis, record NB scores, located in the Primary Organ Systems table, in the appropriate location in Table 2 of this protocol.  Print the ClearView report and export data to NewCalLimitsTesting folder.

  - Repeat step 3.3 for four additional subjects.

  - Calibrate device using brightness setting of 85 and gain setting of 55.  After successful calibration, save images to folder on the network labeled NewCalLimitsTesting.  Load finger images from device validation testing for first subject.  Modify angles as required and submit for analysis.  After analysis, record NB scores in appropriate location in Table 2.  Print the ClearView report and export data to NewCalLimitsTesting folder.

  - Repeat step 3.5 for same four additional subjects.

  - Calibrate device using brightness setting of 100 and gain setting of 63.  After successful calibration, save images to folder on the network labeled NewCalLimitsTesting.  Load finger images from device validation testing for first subject.  Modify angles as required and submit for analysis.  After analysis, record NB scores in appropriate location in Table 2.  Print the ClearView report and export data to NewCalLimitsTesting folder.

  - Repeat step 3.7 for same four additional subjects.

  - Compare original scores to the nominal, low limit, and high limit scores and record Pass in the space provided if score difference is less than 0.5, if more record Fail.

  - Attach printouts of NB portion of labeled ClearView reports to this document.

- Acceptance Criteria

If the NB score values recorded at the three different calibration settings are less than +/- 0.5 from the original NB score values for a subject, the change in calibration pass/fail criteria did not have an impact on the human subject results.  The results and conclusions of this protocol will be summarized into a final report.

Table 1

| Device Serial Number: | Computer System Ref Number: |
| --- | --- |
| Time Device Turned On: | Workstation ID: |
| Exposure Delay: | ClearView Version: |
| Signature: | Date: |

Table 2

Subject 1

| Settings | B85G58 | B85G58 | B85G55 | B85G55 | B100G63 | B100G63 |
| --- | --- | --- | --- | --- | --- | --- |
| Organ System | Physical | Autonomic | Physical | Autonomic | Physical | Autonomic |
| Cardiovascular |  |  |  |  |  |  |
| Respiratory |  |  |  |  |  |  |
| Hepatic |  |  |  |  |  |  |
| Gastrointestinal |  |  |  |  |  |  |
| Renal |  |  |  |  |  |  |
| Pass/Fail |  |  |  |  |  |  |
| Signature |  |  |  |  |  |  |
| Date |  |  |  |  |  |  |

Subject 2

| Settings | B85G58 | B85G58 | B85G55 | B85G55 | B100G63 | B100G63 |
| --- | --- | --- | --- | --- | --- | --- |
| Organ System | Physical | Autonomic | Physical | Autonomic | Physical | Autonomic |
| Cardiovascular |  |  |  |  |  |  |
| Respiratory |  |  |  |  |  |  |
| Hepatic |  |  |  |  |  |  |
| Gastrointestinal |  |  |  |  |  |  |
| Renal |  |  |  |  |  |  |
| Pass/Fail |  |  |  |  |  |  |
| Signature |  |  |  |  |  |  |
| Date |  |  |  |  |  |  |

Subject 3

| Settings | B85G58 | B85G58 | B85G55 | B85G55 | B100G63 | B100G63 |
| --- | --- | --- | --- | --- | --- | --- |
| Organ System | Physical | Autonomic | Physical | Autonomic | Physical | Autonomic |
| Cardiovascular |  |  |  |  |  |  |
| Respiratory |  |  |  |  |  |  |
| Hepatic |  |  |  |  |  |  |
| Gastrointestinal |  |  |  |  |  |  |
| Renal |  |  |  |  |  |  |
| Pass/Fail |  |  |  |  |  |  |
| Signature |  |  |  |  |  |  |
| Date |  |  |  |  |  |  |

Subject 4

| Settings | B85G58 | B85G58 | B85G55 | B85G55 | B100G63 | B100G63 |
| --- | --- | --- | --- | --- | --- | --- |
| Organ System | Physical | Autonomic | Physical | Autonomic | Physical | Autonomic |
| Cardiovascular |  |  |  |  |  |  |
| Respiratory |  |  |  |  |  |  |
| Hepatic |  |  |  |  |  |  |
| Gastrointestinal |  |  |  |  |  |  |
| Renal |  |  |  |  |  |  |
| Pass/Fail |  |  |  |  |  |  |
| Signature |  |  |  |  |  |  |
| Date |  |  |  |  |  |  |

Subject 5

| Settings | B85G58 | B85G58 | B85G55 | B85G55 | B100G63 | B100G63 |
| --- | --- | --- | --- | --- | --- | --- |
| Organ System | Physical | Autonomic | Physical | Autonomic | Physical | Autonomic |
| Cardiovascular |  |  |  |  |  |  |
| Respiratory |  |  |  |  |  |  |
| Hepatic |  |  |  |  |  |  |
| Gastrointestinal |  |  |  |  |  |  |
| Renal |  |  |  |  |  |  |
| Pass/Fail |  |  |  |  |  |  |
| Signature |  |  |  |  |  |  |
| Date |  |  |  |  |  |  |