Approval Signatures:

Author:

| Department | Name | Signature | Date |
| --- | --- | --- | --- |
| Author | Andrea Miller |  |  |

Approval:

| Engineering | Lloyd Kurth |  |  |
| --- | --- | --- | --- |
| OPS | Mike Stowell |  |  |
| Exec | Nancy Rizzo |  |  |
| IS | Andrew Mason |  |  |
| QA | Scott Pletzer |  |  |

- Purpose

This report summarizes the testing performed to validate the Application functions in the ClearView software solution.

- Executive Summary

The EPIC ClearView software Application functions provide the end user with the ability to control user logins and access device settings.  The modification to Version 1.1.1.3 corrected the test duration feature of the Device Testing tab, which previously was using the .5 second duration for all of the settings.

- Protocol Execution and Results

All validation activities were conducted by EPIC staff between 5/2/12 and 5/8/12.  The technician completed the protocol, recording the results directly on the protocol.  The original protocol is located in Attachment A.

  - Protocol Documentation - The protocol required that ClearView Reports be printed.  The hard copy printouts are located in Attachment B of this report.

  - Protocol Electronic Files - The protocol required that data be saved for certain verification tests.    The CD of this recorded data is located in Attachment C.

  - Deviations - deviations from the protocol were recorded in Table 1 below.

Table 1

Deviations

| No. | Description of Non-conformance (include reference to the protocol section) | Resolution/Action Taken |
| --- | --- | --- |
| 1 | Page 8, Section 6.3.2 – “aclearview” user does not exist, users beginning with “aclearview” do exist. | The script was modified to select any of the “aclearview” users as the test remains the same. The next step was removed because it would duplicate. The protocol was updated to modify instructions to match those executed here. |

  - Non-conformances - non-conformances from the protocol were recorded in Table 2 below.

Table 2

Non-Conformances

| No. | Description of Non-conformance (include reference to the protocol section) | Resolution/Action Taken |
| --- | --- | --- |
| 1 | Page 17, Section 6.3.4 – The images taken during the duty cycle test at Exposure Delay 150 did not display the probe, the images were either black or energized. | This test is intended to determine that changes made to the exposure delay button alter the device settings. Although the LED light was not lit, the black image represents that the timing was changed (exposure delay) to a point before the plate is energized. The protocol will be modified to either ensure the LED light is illuminated and the probe is present or black image (not energized) is present. This test is a confirmation that the change in exposure delay setting results in a change to the image captured. |
| 2 | Page 26, Section 6.5 – Auto calibration did not produce a passing calibration image set, could not proceed. Device not responding to ClearView commands. Retest - attempt #1 to auto calibrate the device resulted in a connectivity error. All 10 images were black/blank and the device did not beep during calibration. Exited out of ClearView and turned the device off and back on, and unplugged the USB cable. During attempt #2, the device took one warm-up image then lost connectivity and the remaining images were blank. Exited out of ClearView and restarted the computer. Unplugged the USB cable to verify connectivity of the device before starting ClearView. | The test technician modified the device settings until a successful calibration was obtained. Once the settings were established, the technician retested the entire section starting where the failure occurred. Upon retest, all validation requirements were met. |

One recorded non-conformance demonstrated repeated difficulty in maintaining the communication between the device and software.  These errors occur when the USB connection between the device and the software is interrupted.  The error was resolved in each instance by the technician.  The resolution requires the technician to shut down the software and/or the computer system, restarting either the application and/or both the application and the computer system.  Upon successful initializing of the ClearView software, each error was resolved and the software was found to perform as expected.

These communication failures have been categorized as anomalies since the technician was able to demonstrate that the software works as expected once the communication interruption is resolved (i.e., mitigation occurs).  Device/software communication anomalies can be mitigated by the user through restarting the application and/or the computer system.  The cause of this anomaly will continue to be investigated.

| Table 1: Unresolved Anomalies |  |  |
| --- | --- | --- |
| Anomaly | Impact on Performance | Plans for Correction |
| Device not responding to ClearView software commands. | ClearView Software will not collect energized images. User must shut down application and restart computer and collect images. | See Memo - ClearView System Communication Errors dated 5/15/12. |

- Conclusions

The Applications Functions of version 1.1.1.3 of ClearView software are considered validated for use in IDE investigations with the recommendation that users are instructed how to detect and resolve communication errors between the device and the software.

- Attachments

Attachment A - Original Signed Protocol

Attachment B – Validation Test Printouts

Attachment C – CD of Downloaded Files