Approval Signatures:

| Department | Name | Signature | Date |
| --- | --- | --- | --- |
| Author | Dan Miller |  |  |

Approval:

| OPS | Mike Stowell |  |  |
| --- | --- | --- | --- |
| Exec | Nancy Rizzo |  |  |
| IS | Andrew Mason |  |  |
| QA | Scott Pletzer |  |  |

- Purpose

This report summarizes the testing performed to verify and validate the Camera functions in the ClearView software.

- Scope

The EPIC ClearView Software Camera function provides the end user with the ability to perform calibration verification.  In version 1.1.1.3.1, the software was modified to address a usability issue and calibration patch.  A description of the changes made and the rationale behind the construction of this protocol are below.

The reason for the software patch was the realization that when the camera in the scanner was first initialized via USB it would not accurately set the Gain value to the value specified in the configuration. This would only happen on the first connection to the camera and never on subsequent connections. This has since been attributed to a problem with the initialization of the camera driver. The change to the code prevents the initialization problem by rearranging the calls that are made to set the appropriate settings on the camera and to connect to the video stream. The process will now connect the video stream before sending the desired settings. The change only alters the order commands that are issued and does not impact any other functionality.

The second change made was to modify the acceptable range of pixel intensity values within the matrix created by a calibration image.  The pixel intensity matrix focuses on three regions of importance within the image. There is the inner circle (dark space), the middle ring which consists of the brightest ring that degrades with outward radial distance and the outer ring that falls off into the dark space around the image. The new camera design produces images that match the 1st and 2nd ring areas but the third ring is typically smaller by a margin only slightly larger than what has historically been seen. The current matrix has the three limits of acceptable levels – inner circle area is set to +/- 3 Standard deviation, the middle ring (most critical) is set to +/- 1 standard deviation, and the outer ring is set to +/- 3 standard deviations. The only change to this process will be to allow the outer ring to be +3/-4 standard deviations.

- Executive Summary

- Protocol Execution and Results

All validation activities were conducted by EPIC staff on 7/9/12, 7/10/12 and 7/18/12.  The technician completed the protocol, recording the results directly on the protocol.  The original protocol is located in Attachment A.  A CD of electronic files generated during the validation testing is located in Attachment B.  All ClearView printouts generated during the execution of the protocol are located in Attachment C.

  - Deviations - Deviations were recorded in Table 1 below.

Table 1

Deviations

| No. | Description of Deviation (include reference to the protocol section) | Resolution/Action Taken |
| --- | --- | --- |
| 1 | Page 17, Section 6.3.6 – Changed the exposure delay setting from 200 to 290 because 290 would not capture all 10 images successfully. | Faint and blank images are expected at the specified setting (200). Technician interpreted protocol to capture 10 visible images. Protocol will be modified to clarify this. Retest where deviation was made and proceed thru page 18. Any non-conformances will be recorded as specified in the protocol. |

  - Non-conformances - non-conformances were recorded in Table 2 below.

Table 2

Non-Conformances

| No. | Description of Non-conformance (include reference to the protocol section) | Resolution/Action Taken |
| --- | --- | --- |
| 1 | Page 16, Section 6.3.6 – Calibration was supposed to fail, but it passed with an improperly seated probe. | Protocol instructions were modified to accomplish improper seating of probe. Retest starting at 6.3.6 and proceed thru the middle of page 17. Non-conformances recorded on pages 19 and 21 are due to the calibration in the original protocol execution passing and therefore failure data did not appear as expected in the View Alerts tab. |

- Conclusions

The Camera Functions of version 1.1.1.3.1 of ClearView Software Solution are considered validated for use in IDE investigations.

- Attachments

Attachment A- Original Signed Protocol

Attachment B- CD of electronic files

Attachment C- ClearView Camera Validation Printouts