Approval Signatures:

| Department | Name | Signature | Date |
| --- | --- | --- | --- |
| Author | Scott Pletzer |  |  |

Approval:

| OPS | Mike Stowell |  |  |
| --- | --- | --- | --- |
| Exec | Nancy Rizzo |  |  |
| IS | Andrew Mason |  |  |
| MFG/Indep. | Dan Miller |  |  |

- Purpose

This report summarizes the testing performed to verify and validate the functionality of EPIC ClearView software version 1.1.1.4, in order to release it for use in conjunction with the requirements of DDP-008.000, “Design & Development Plan (DDP) ClearView 1.1.1.4”.

- Scope

Prior to the testing of software version 1.1.1.4, the software was modified to version 1.1.1.3.1 (Patch), to address a usability issue and calibration patch, but was never released for use.  The verification and validation testing of EPIC ClearView Software version 1.1.1.3.1 (Patch) was documented under ENG-010.005, “Camera Functions ClearView Software v. 1.1.1.3.1 (Patch) Verification and Validation Report”.  The rationale and justification for changing the calibration acceptance limits with the software was documented under ENG-030.000, “Qualifying the Calibration Acceptance Limits Change Final Report”.

Prior to the release of version 1.1.1.3.1 (Patch), the software was modified to improve the design and performance in the following areas:

- Calibration

- Camera Setting Persistence

- Scanner Connectivity

This report documents the verification and validation testing performed in order to confirm the functionality of version 1.1.1.4 so that it may be released for use.

- Executive Summary

The verification and validation testing of version 1.1.1.4 was documented by the execution of the following three protocols:

- ENG-009.006, ClearView Software Validation Protocol – APPLICATION Functions

- ENG-010.006, ClearView Software Validation Protocol – CAMERA Functions

- ENG-011.005, ClearView Software Validation Protocol – CAPTURE Functions

These executed protocols are contained in Appendices A, B, and C respectively.  Any deviations or nonconformances associated with these protocols are described in detail in the section below.

- Protocol Execution and Results

All validation activities were conducted by EPIC staff on 8/2/12 through 8/7/12.  The technicians completed the protocols, recording the results directly on the protocols.  Any associated deviations or nonconformances are summarized below and their resolutions are documented in the attached protocols.  The original protocols are located in Appendices A, B, and C.  A CD of electronic files generated during the validation testing is located in Appendix D.

Table 1

Protocol Execution Results

| Protocol No. | Test Dates | Results of protocol | Original Protocol Attached in: |
| --- | --- | --- | --- |
| ENG-009.006 | 8/7/12 | Deviations - NONE Nonconformances - NONE | Appendix A All original printouts generated during protocol execution are located behind the protocol in Attachment A. |
| ENG-010.006 | 8/2/12 through 8/7/12 | Deviations - NONE Nonconformances - Page 18, only one alert appeared. - Page 20, only two alerts appeared. | Appendix B All original printouts generated during protocol execution are located behind the protocol in Attachment B. |
| ENG-011.005 | 8/6/12 through 8/7/12 | Deviations - NONE Nonconformances - Page 24, no raw data is displayed in Admin tab. Also, the NS values are not displayed (Also occurs on page 37, 59, 62, 63, and 69) | Appendix C All original printouts generated during protocol execution are located behind the protocol in Attachment C. |

- Conclusions

EPIC ClearView software version 1.1.1.4 is considered verified and validated for use, in conjunction with the requirements DDP-008.000.

- Attachments

Appendix A - Original Signed Protocol (ENG-009.006)

Appendix B - Original Signed Protocol (ENG-010.006)

Appendix C - Original Signed Protocol (ENG-011.005)

Appendix D - CD of Electronic Files