Approval Signatures:

| Department | Name | Signature | Date |
| --- | --- | --- | --- |
| Author | Scott Pletzer |  |  |

Approval:

| OPS | Mike Stowell |  |  |
| --- | --- | --- | --- |
| Exec | Nancy Rizzo |  |  |
| IS | Andrew Mason |  |  |
| MFG/Indep. | Dan Miller |  |  |

Revision History

| Revision | Revision Date | Reason for Revision/Change Request | Revised By |
| --- | --- | --- | --- |
| 007 | 25Sep12 | Revised to accommodate the testing requirements for the release of version 2.0. | S. Pletzer |

- Purpose

This report summarizes the testing performed to verify and validate the functionality of EPIC ClearView software version 2.0, in order to release it for use in conjunction with the requirements of DDP-007.000, “Design & Development Plan (DDP) ClearView 2.0”.

- Scope

This new release of ClearView is required because the existing iteration of the product has a significant amount of unused (also referred to as dead) code that needed to be removed in order to improve maintainability.  In addition there were over 120 database tables that were no longer used that were removed, pre-scan questions were included, and a new report interface was introduced to display the analysis results in a more easy to understand manner.

This report documents the verification and validation testing performed in order to confirm the functionality of version 2.0 so that it may be released for use.

- Executive Summary

The verification and validation testing of version 2.0 was documented by the execution of the following eight protocols:

- ENG-009.007, ClearView Software Validation Protocol – APPLICATION Functions

- ENG-009.008, ClearView Software Validation Protocol – APPLICATION Functions

- ENG-010.007, ClearView Software Validation Protocol – CAMERA Functions

- ENG-010.008, ClearView Software Validation Protocol – CAMERA Functions

- ENG-011.006, ClearView Software Validation Protocol – CAPTURE Functions

- ENG-011.007, ClearView Software Validation Protocol – CAPTURE Functions

- ENG-011.008, ClearView Software Validation Protocol – CAPTURE Functions

- ENG-012.002, ClearView Software Validation Protocol – PATIENT DEMOGRAPHICS Functions

- ENG-014.005, ClearView Software Validation Protocol – REPORT Functions

These executed protocols are contained in Appendices A, B, C, D, E, F, G, H, and I respectively.  Any deviations or nonconformances associated with these protocols are described in detail in the section below.

- Protocol Execution and Results

All validation activities were conducted by EPIC staff between 8/30/12 and 9/25/12.  The technicians completed the protocols, recording the results directly on the protocols.  Any associated deviations or nonconformances are summarized below and their resolutions are documented in the attached protocols.  If more than two revisions of the same protocol are listed below, a deviation and/or nonconformance drove the need to repeat a portion of the protocol to satisfy requirements.  The one exception was ENG-011 which was revised a second time to revision 008.  This was done after a review of the documentation revealed that not all portions of the protocol were properly changed and executed to satisfy the nonconformances.  The original protocols are located in Appendices A-I, and a CD of electronic files generated during the validation testing is located in Appendix J.

Table 1

Protocol Execution Results

| Protocol No. | Test Dates | Results of protocol | Original Protocol Attached in: |
| --- | --- | --- | --- |
| ENG-009.007 | 8/30/12 through 8/31/12 | Deviations – NONE Nonconformances - Section 6.3.4, page 10 – Brightness in the camera properties tab does not match the image in the protocol of the camera properties. - Section 6.3.4, page 11 – The images taken in Duty Cycle Testing do not appear to be dimmer even after the setting was changed to a low gain. - Section 6.3.4, page 12 – The gain slider was all the way to the right in camera properties even though it should be more in the middle of the bar. - Section 6.4, page 22 – Settings icon is not displayed when logged in as a Tech Administrator. - Section 6.4, page 23 – After clicking New Capture, the pre-scan questions appear instead of the New Capture screen. | Appendix A All original printouts generated during protocol execution are located behind the protocol in Attachment A. |
| ENG-009.008 | 9/4/12 | Deviations - NONE Nonconformances - NONE | Appendix B All original printouts generated during protocol execution are located behind the protocol in Attachment B. |
| ENG-010.007 | 9/5/12 | Deviations - The Camera Functions test is being done on a cart configuration that does not use a UPS or printer (the device, etc. is plugged into the wall). Nonconformances - Section 6.3.6, page 17 – After clicking “OK” to save the image to the network, an error message appears that says “unable to save images”. - Section 6.3.6, page 21 – No image was saved in the folder. - Section 6.4, page 28 – Only 9 images saved into the folder, not all 10. | Appendix C All original printouts generated during protocol execution are located behind the protocol in Attachment C. |
| ENG-010.008 | 9/6/12 | Deviations - NONE Nonconformances - NONE | Appendix D All original printouts generated during protocol execution are located behind the protocol in Attachment D. |
| ENG-011.006 | 9/7/12 through 9/11/12 | Deviations - Section 6.4, page 18 – Need to add a step fo capturing an image for 2R on the images with capacitive barrier row. Nonconformances - Section 9.1.2, pages 56/57 – After clicking “New Capture”, the pre-scan questions appeared before the New Capture screen. - Section 9.1.12, page 69 – The finger image is not displayed after turning the device back on and clinking New Capture.. | Appendix E All original printouts generated during protocol execution are located behind the protocol in Attachment E. |
| ENG-011.007 | 9/12/12 | Deviations - NONE Nonconformances - NONE | Appendix F All original printouts generated during protocol execution are located behind the protocol in Attachment F. |
| ENG-011.008 | 9/25/12 | Deviations - NONE Nonconformances - NONE NOTE - A review of the documentation revealed that not all portions of revision 007 of the protocol were properly changed and executed to satisfy the nonconformances found during the execution of revision 006. | Appendix G All original printouts generated during protocol execution are located behind the protocol in Attachment G. |
| ENG-012.002 | 8/30/12 | Deviations - NONE Nonconformances - NONE | Appendix H All original printouts generated during protocol execution are located behind the protocol in Attachment H. |
| ENG-014.005 | 9/13/12 through 9/14/12 | Deviations – NONE Nonconformances - Section 6.3.8, page 25 – The cardiovascular system needs the heart to be labeled differently after the review by Nancy Rizzo. Also, the small intestine section needs “Blind Gut” to be labeled “Blind Gut / Stomach”. | Appendix I All original printouts generated during protocol execution are located behind the protocol in Attachment I. |

- Conclusions

EPIC ClearView software version 2.0 is considered verified and validated for use, in conjunction with the requirements DDP-007.000.  The use of version 2.0 is restricted for internal development purposes only, and not for clinical applications.

- Attachments

Appendix A - Original Signed Protocol (ENG-009.007)

Appendix B - Original Signed Protocol (ENG-009.008)

Appendix C - Original Signed Protocol (ENG-010.007)

Appendix D - Original Signed Protocol (ENG-010.008)

Appendix E - Original Signed Protocol (ENG-011.006)

Appendix F - Original Signed Protocol (ENG-011.007)

Appendix G - Original Signed Protocol (ENG-011.008)

Appendix H - Original Signed Protocol (ENG-012.002)

Appendix I - Original Signed Protocol (ENG-014.005)

Appendix J - CD of Electronic Files