Approval Signatures:

| Department | Name | Signature | Date |
| --- | --- | --- | --- |
| Author | Scott Pletzer |  |  |

Approval:

| OPS | Mike Stowell |  |  |
| --- | --- | --- | --- |
| Exec | Nancy Rizzo |  |  |
| IT | Andrew Mason |  |  |
| IT | Lloyd Kurth |  |  |
| MFG/Indep. | Dan Miller |  |  |

Revision History

| Revision | Revision Date | Reason for Revision/Change Request | Revised By |
| --- | --- | --- | --- |
| 007 | 25Sep12 | Revised to accommodate the testing requirements for the release of version 2.0. | S. Pletzer |
| 008 | 16Nov12 | Revised to accommodate the testing requirements for the release of version 2.1. | S. Pletzer |

- Purpose

This report summarizes the testing performed to verify and validate the functionality of EPIC ClearView software version 2.1, in order to release it for use in conjunction with the requirements of DDP-009.000, “Design & Development Plan (DDP) ClearView 2.1”.

- Scope

This new release of ClearView is required because the calibration routine used in ClearView is being updated to use a newer more encompassing algorithm, without a change to the base algorithm.  This update will yield a more robust calibration routine that looks at the images in five distinct areas rather than the current process of two areas.  The current calibration routine although good has limited sensitivity in specific areas, where the new routine is more sensitive.

This report documents the verification and validation testing performed in order to confirm the functionality of version 2.1 so that it may be released for use.

- Executive Summary

The verification and validation testing of version 2.1 was documented by the execution of the following eight protocols:

- ENG-009.009, ClearView Software Validation Protocol – APPLICATION Functions

- ENG-009.010, ClearView Software Validation Protocol – APPLICATION Functions

- ENG-010.009, ClearView Software Validation Protocol – CAMERA Functions

- ENG-011.009, ClearView Software Validation Protocol – CAPTURE Functions

- ENG-011.010, ClearView Software Validation Protocol – CAPTURE Functions

- ENG-012.003, ClearView Software Validation Protocol – PATIENT DEMOGRAPHICS Functions

- ENG-014.006, ClearView Software Validation Protocol – REPORT Functions

- ENG-037.000, ClearView Software Validation Protocol – CALCULATION Functions

These executed protocols are contained in Appendices A, B, C, D, E, F, G, and H respectively.  Any deviations or nonconformances associated with these protocols are described in detail in the section below.

- Protocol Execution and Results

All validation activities were conducted by EPIC staff between 10/24/12 and 11/16/12.  The technicians completed the protocols, recording the results directly on the protocols.  Any associated deviations or nonconformances are summarized below and their resolutions are documented in the attached protocols.  If more than two revisions of the same protocol are listed below, a deviation and/or nonconformance drove the need to repeat a portion of the protocol to satisfy requirements.  The original protocols are located in Appendices A-H, and a CD of electronic files generated during the validation testing is located in Appendix I.  All original printouts generated during each protocol execution are located behind the respective protocol.

Table 1

Protocol Execution Results

| Protocol No. | Test Dates | Results of protocol | Original Protocol Location: |
| --- | --- | --- | --- |
| ENG-009.009 | 10/25/12 | Deviations – NONE Nonconformances - Section 6.5, page 25 – When logged in as an Admin., unable to save the modified patient demographics. | Appendix A |
| ENG-009.010 | 11/1/12 through 11/2/12 | Deviations - NONE Nonconformances - NONE | Appendix B |
| ENG-010.009 | 11/2/12 through 11/5/12 | Deviations - NONE Nonconformances - NONE | Appendix C |
| ENG-011.009 | 10/26/12 | Deviations - Section 8.1, page 55 – Under the Alerts tab, the message section needs to reflect the new cal. Failure verbiage. Nonconformances - Section 8.1, page 56 – Setting “Configuration 2” does not pass the calibration routine. | Appendix D |
| ENG-011.010 | 11/2/12 | Deviations - NONE Nonconformances - NONE | Appendix E |
| ENG-012.003 | 10/24/12 | Deviations - NONE Nonconformances - NONE | Appendix F |
| ENG-014.006 | 10/24/12 through 10/25/12 | Deviations – NONE Nonconformances - NONE | Appendix G |
| ENG-037.000 | 10/29/12 through 11/16/12 | Deviations – NONE Nonconformances - 3 of the 498 data points checked were not exact. For subject 1, the left filtered nervous system value was reporting as 14 rather than 15. For subject 2, the unfiltered NB score for respiratory was reporting as 1.1 rather than 1. For subject 3, the unfiltered NB respiratory value was reporting as 1.4 rather than 1.3. | Appendix H |

- Conclusions

EPIC ClearView software version 2.1 is considered verified and validated for use, in conjunction with the requirements DDP-009.000.

- Attachments

Appendix A - Original Signed Protocol (ENG-009.009)

Appendix B - Original Signed Protocol (ENG-009.010)

Appendix C - Original Signed Protocol (ENG-010.009)

Appendix D - Original Signed Protocol (ENG-011.009)

Appendix E - Original Signed Protocol (ENG-011.010)

Appendix F - Original Signed Protocol (ENG-012.003)

Appendix G - Original Signed Protocol (ENG-014.006)

Appendix H – Original Signed Protocol (ENG-037.000)

Appendix I - CD of Electronic Files