Approval Signatures:

Author:

| Department | Name | Signature | Date |
| --- | --- | --- | --- |
| Author | Andrea Miller |  |  |

Approval:

| Engineering | Lloyd Kurth |  |  |
| --- | --- | --- | --- |
| OPS | Mike Stowell |  |  |
| Exec | Nancy Rizzo |  |  |
| IS | Andrew Mason |  |  |
| QA | Scott Pletzer |  |  |

- Purpose

This report summarizes the testing performed to validate the Capture functions in the ClearView software.

- Executive Summary

The EPIC ClearView software Capture functions provides the end user with a method for adding new patients to the database, searching patient database, and perform a new capture/measurement of a patient using the ClearView hardware device.  In version 1.1.1.3, a modification was made to the timing mechanism used to capture the image in the calibration process and the capture process in the ClearView software.  This change was to code design only and did not modify the image capture timing in any manner.

- Protocol Execution and Results

All validation activities were conducted by EPIC staff between 5/4/12 and 5/9/12.  The results were recorded directly on the protocol.  This original protocol is located in Attachment A.

  - Protocol Outputs- The protocol required that ClearView Reports be printed.  The hard copy printouts are located in Attachment B of this report.

  - Protocol Electronic Files - The protocol required that data be saved for certain verification tests.    The CD of this recorded data is located in Attachment C.

  - Deviations- No deviations were recorded during the execution of this protocol.

  - Non-conformance- Table 1 below summarizes the non-conformances that occurred during testing and the resolution/action taken for each.

| Table 1 Non-Conformances |  |  |
| --- | --- | --- |
| No. | Description of Non-conformance (include reference to the protocol section) | Resolution/Action Taken |
| 1 | Page 64, Section 9.1.7 – Blue screen (BSOD) occurred during patient image capture. Computer and scanner were restarted. | These two non-conformances were handled through complete retest starting at Section 9.1.6 through the remaining portions of the protocol. Upon retest, the image collection and analysis functions performed as expected. |
| 2 | Page 65, Section 9.1.8 – The ClearView software froze and became non-responsive after the first patient image capture. Computer and scanner were restarted. | See #1 above. |
| 3 | Page 62, Section 9.1.5 – The finger box did not have a finger image in it. | The technician presented a subject which has finger images and re-executed Section 9.1.5 (all steps). Upon retest, the image collection and analysis functions performed as expected. |
| 4 | Page 62, Section 9.1.5 – The software does not allow a patient image set to be loaded with the live finger, only the energized image can be loaded. | Requirement 4.48 of the Capture Functions Requirements document (SR-202) specifies that the lit finger image will not be loaded during the administrator loading processes. The software works as expected and the protocol will be updated to remove the requirement to load lit finger images. |

Two non-conformances were recorded for which a cause could not be identified (i.e., #1 and #2).  Both recorded non-conformances demonstrated difficulty in maintaining the communication between the device and software.  These errors can occur when the USB connection between the device and the software is interrupted.  The error was resolved in each instance by the technician.  The resolution required the technician to shut down the software and the computer system, restarting both the application and the computer system.  Upon initializing the ClearView software, each error was resolved and the software was found to perform as expected.

These communication failures are being categorized as anomalies since the technician was able to demonstrate that the software works as expected once the communication interruption is resolved (i.e., mitigation occurred).  Device/software communication anomalies can be mitigated by the user through restarting the application and/or the computer system.  The cause of this anomaly is under investigation.

| Table 1: Unresolved Anomalies |  |  |
| --- | --- | --- |
| Anomaly | Impact on Performance | Plans for Correction |
| Received “Unable to Perform Image Calculations” error and a subsequent Blue Screen of Death (BSOD). | The user attempts to resolve the error by clicking “OK” and the computer operating system spontaneously launches a BSOD, the application and the system is shut down. The user loses functionality for the period of the shut down and the restart. User can begin to use the system after restart successfully. | See Memo - ClearView System Communication Errors dated 5/15/12. |
| Device not responding to ClearView software commands. | ClearView Software will not collect energized images. User must shut down application and restart computer and collect images. |  |

Other than the two unresolved anomalies listed above, the Capture Functions were found to perform as expected for all requirements.

- Conclusions

The Capture Functions of version 1.1.1.3 of ClearView software are considered validated for use in IDE investigations with the recommendation that users are instructed how to detect and resolve communication errors between the device and the software.

- Attachments

Attachment A- Original Signed Protocol

Attachment B- ClearView Report Validation Printouts

Attachment C – CD of Downloaded Files