This document was developed by NAMSA personnel, an EPIC consultant.  EPIC’s engineering study control number was assigned to the document for use in referencing in other EPIC documentation.  Version control from NAMSA is indicated in the footer using the “v1” designation.

This engineering study was assigned the control number ENG-025 revision 000 on 2/10/12.

EPIC ClearViewTM

Algorithm Development

Version 1.0

EPIC Research & Diagnostics, Inc.

8501 E Princess Dr. Suite 100

Scottsdale, AZ  85255

USA

(480) 477-5242 Phone

(480) 292-7933 Fax

EPIC ClearView Algorithm Development Report

Approved By:

Jennifer Mischke, MPH						Date

Principal Consultant, Biostatistics

NAMSA

Nancy Rizzo, PhD						Date

Founder and CEO

EPIC Research & Diagnostics

Andrew Mason							Date

CTO

EPIC Research & Diagnostics

Andrea Miller							Date

Quality Assurance Manager

EPIC Research & Diagnostics

1	Introduction	5

2	Algorithm Development Data	5

2.1	Development Subjects	6

2.2	ClearView Scan Coefficients	7

2.3	Data Transformations	8

3	Statistical Methods	9

3.1	Modeling Procedures	9

3.1.1	Logistic Regression	10

3.1.2	Naïve Bayes Classifier	10

3.1.3	Ridge Regression, LASSO, and LARS	11

3.2	Cross-Validation	11

3.3	Model Evaluation	12

4	Calculation of the ClearView Response Scale	13

4.1	Epic Score Validator	13

4.2	Sector Scores	13

4.2.1	Logistic Regression Scoring	13

4.2.2	Sector Scoring	13

4.3	Primary Scores	14

5	Validation of Implementation	15

6	Appendix A: Description of ClearView Scan Coefficients	16

6.1	Sector Area and Normalized Area	16

6.2	Average Intensity	16

6.3	Entropy	16

6.4	Form	16

6.5	Form 2 and Form2’	16

6.6	Fractal	17

6.7	NS Parameter	17

7	Appendix B: Mapping 78 Sectors to 47 Organs	18

8	Appendix C: Subjects Selected for SAS vs ClearView Verification	21

# Introduction

The EPIC ClearView is a galvanic skin response (GSR) measurement tool for the evaluation of the human body’s 5 primary organ systems, including: 1) Cardiovascular, 2) Respiratory, 3) Gastrointestinal and Endocrine, 4) Hepatic,5) Renal and Reproductive.

The characterization of the ClearView GSR is captured by the Response Scale. The Response Scale is a score, ranging from 0-5 for each of the five primary organ systems above. Values that fall higher on the scale are intended to indicate an increased level of association or concern with a diagnosis.

The purpose of this document is to describe the statistical development of the EPIC ClearView Response Scale. While years of working with the ClearView has laid the foundation for understanding the relationship between the scan results and the health of the individual, extensive statistical modeling was conducted to finalize the algorithm and Response Scale.

This document describes the development (training) data used for the statistical modeling, as well as the methods used.

# Algorithm Development Data

To begin, this section describes the subject data used to develop the EPIC ClearView Response Scale.  For the purpose of this document, the data used for algorithm development are referred to interchangeably as the ‘development cohort’ or the ‘training dataset.’

The development cohort consists of 345 subjects from 5 separate investigations, which are described more fully in Section 2.1.

The 345 subjects yielded a large volume of data. Each subject underwent 2 sets of ClearView scans: 1 without the filter and 1 with the filter in place. Each scan included each of the 10 fingers. The 10 fingers were divided into a total of 78 sectors, with 19 scan coefficients measured for each sector. Therefore, a single subject produced 2,964 separate measures (19 scan coefficients x 78 sectors x 2 sets of scans).

For the purpose of statistical modeling, the filtered and not filtered sets of ClearView scans were treated separately.

The health status of each subject was collected with respect to 5 separate organ or system groups: cardiovascular system, gastrointestinal/endocrine system, respiratory system, renal system, and the hepatic system. These 5 systems were considered independently; each subject was given a yes/no (1/0) value for each of the 5 organ or systems groups, based upon absence or presence of pre-identified qualifying diagnoses within each group. For the purpose of this document, the 5 organ or systems groups will be referred to as “diagnosis groups.” If a subject is free from all qualifying diagnoses across the 5 groups, that subject is defined as a control.

## Development Subjects

The dataset used for algorithm development consisted of 345 subjects from 5 separate studies, which are briefly described in Table 1. For the purpose of algorithm development, subjects were included that covered a range of geographies, ages, and health states. This was an intentional mix of subjects; the data used for algorithm development was intended to match the intended population of use.

Table 1. Data Sources for the Algorithm Development Cohort

| Study Name | Time Frame | Geographic Area | N | Description of Subjects | Basic Study Design |
| --- | --- | --- | --- | --- | --- |
| GBMC | SEP2010 – JUN2011 | In-patient and outpatient patients at the Greater Baltimore Medical Center in Baltimore, Maryland. | 83 | Men and women ages 18-85 with at least one qualifying diagnosis (case). | Two ClearView scans during a single study visit; compare to medical diagnoses. |
| DV-001 | MAR2011 – JUN2011 | Friends, family, and staff of EPIC Research and Diagnostics, Inc. at EPIC’s corporate office in Scottsdale, Arizona. | 92 | Men and women ages 18-73 with varying health conditions. Most subjects were free from “qualifying” diagnoses as categorized in the WebForm. | One set of ClearView scans during a single study visit; compare to medical diagnoses. |
| DV-002 | MAR2011 – MAY2011 | Friends, family, and staff of EPIC Research and Diagnostics, Inc. at EPIC’s corporate office in Scottsdale, Arizona. | 19 | Men and women ages 24-64 with varying health conditions. Most subjects were free from “qualifying” diagnoses as categorized in the WebForm. | One ClearView scan per week for at least 7 out of 8 consecutive weeks on the same day at the same time each week. Most subjects had 8 visits. |
| HVS | JUN2010 | Patients of Southwest Surgery and Midwest Internal Medicine in Lake Havasu City, Arizona. | 52 | Men and women ages 34-88 with a wide array of conditions, including many “qualifying” diagnoses as categorized in the WebForm. | One ClearView scan during a single study visit; compare to medical diagnoses. |
| EHI | JUN2005 – APR2008 | Patients of the Enlightened Health Institute in Scottsdale, Arizona. | 99 | Men and women ages 16-76 with a wide array of conditions. | One ClearView scan during a single study visit; compare to medical diagnoses. |

The subjects used for algorithm development comprise data unique from the GBMC clinical study. That is, subjects used for development were excluded from the analysis of the GBMC clinical trial so that the assessment of the performance of the ClearView Response Scale could be validated on an independent dataset representative of the population of intended use. To maintain this clear separation of data, a randomly selected subset of subjects from the GBMC study was used for algorithm development and was excluded from the analyses of the clinical trial. For the four studies other than GBMC, all available subjects were used for the algorithm development cohort.

The subjects in the training dataset were 49% male, with an average age of 53 ± 14.6 years. The development cohort included subjects with a range of medical diagnoses, as well as subjects free from an active medical issue (to serve as controls), as described in Table 2.

Table 2. Demographic Characteristics for the Algorithm Development Cohort

| Characteristic1 | DV001 (N = 92) | DV002 (N = 19) | EHI (N = 99) | GBMC Subset3 (N = 83) | HVS (N = 52) | Overall (N = 345) |
| --- | --- | --- | --- | --- | --- | --- |
| Age (years) Mean ± SD Median (Min, Max) | 47.6 ± 13.6 51.5 (19, 73) | 43.9 ± 14.6 44.0 (25, 65) | 50.9 ± 12.9 53.0 (17, 77) | 56.3 ± 14.0 57.0 (21, 88) | 64.8 ± 11.6 65.0 (35, 89) | 53.0 ± 14.6 54.0 (17, 89) |
| Gender: Male | 56.5% (52/92) | 52.6% (10/19) | 36.4% (36/99) | 54.2% (45/83) | 48.1% (25/52) | 48.7% (168/345) |
| Qualifying Diagnosis2 Cardiovascular GI/Endocrine Renal Hepatic Respiratory Control Group | 14% (13/92) 3% (3/92) 0% (0/92) 0% (0/92) 2% (2/92) 84% (77/92) | 11% (2/19) 0% (0/19) 0% (0/19) 0% (0/19) 5% (1/19) 84% (16/19) | 4% (4/99) 1% (1/99) 0% (0/99) 0% (0/99) 1% (1/99) 96% (95/99) | 66% (55/83) 45% (37/83) 28% (23/83) 18% (15/83) 31% (26/83) 2% (2/83) | 56% (29/52) 17% (9/52) 0% (0/52) 2% (1/52) 8% (4/52) 40% (21/51) | 30% (103/345) 14% (50/345) 7% (23/345) 5% (16/345) 10% (34/345) 61% (211/345) |

1 Continuous measures are presented as Average ± Standard Deviation; Median; (Minimum, Maximum).

Categorical measures are presented as Proportion (x/n).

2 Qualifying Diagnosis Groups are not mutually exclusive, so that the sum of percentages could be greater than 100%.

3 The GBMC Development Subset consists of 83 randomly selected participants whose study data was set aside and used for algorithm

development only.

## ClearView Scan Coefficients

The ClearView scan coefficients are characterizations of the physical attributes of the scan images. When a scan is conducted, the images from all 10 fingers are divided into 78 sectors. Each sector is then described using the 19 ClearView scan coefficients.

The scan coefficients are listed in Table 3, and are briefly described in Appendix A.

Table 3. ClearView Scan Coefficients

ClearView Scan Coefficient

Average Intensity

Average Intensity (1)

Average Intensity (2)

Average Intensity (3)

Average Intensity (4)

Break

Entropy

Fractal

Form

Form (1)

Form (2)

Form (3)

Form (4)

Form 2

Form 2 Prime

NS Integer

Normalized Area

Ring Intensity

Ring Thickness

## Data Transformations

Each of the individual scan coefficients was assessed for normality. If the distribution appeared skewed or there was evidence of kurtosis, a log transformation was implemented. For scan coefficients that had the potential of negative values, a value of 5 was added prior to taking the log. Additionally, one coefficient was dichotomized; the scan coefficient referred to as “BREAK” was given a value of 1 if the original value was non-zero, and a value for 0 if the original value was also 0. This dichotomization of the BREAK scan coefficient was because most values were 0, with non-zero values covering a very wide range with no clear distribution.

Table 4 defines any transformations conducted on the ClearView scan coefficients, as well as the calculation of age and gender.

Table 4. Transformations of ClearView Scan Coefficients

| ClearView Scan Coefficient | Variable Name | Data Transformation |
| --- | --- | --- |
| Age | AGE | (SCAN_DATE - DOB)/365.25 |
| Sex | MALE | MALE = 1 if SEX = ‘M’ MALE = 0 if SEX = ‘F’ |
| Average Intensity | AI | -- |
| Average Intensity (1) | AI1 | -- |
| Average Intensity (2) | AI2 | -- |
| Average Intensity (3) | AI3 | LOG(AI3+5) |
| Average Intensity (4) | AI4 | LOG(AI4 + 5) |
| Break | BREAKCOEFFICIENT | if BREAKCOEFFICIENT = 0 then BREAKBIN = 0 else if BREAKCOEFFICIENT > 0 then BREAKBIN = 1 |
| Entropy | ENTROPY | -- |
| Fractal | FRACTALCOEFFICIENT | -- |
| Form | FORM1 | -- |
| Form (1) | FORM1_1 | LOG(FORM1_1) |
| Form (2) | FORM1_2 | LOG(FORM1_2) |
| Form (3) | FORM1_3 | LOG(FORM1_3 + 5) |
| Form (4) | FORM1_4 | LOG(FORM1_4 + 5) |
| Form 2 | FORM2 | LOG(FORM2) |
| Form 2 Prime | FORM2PRIME | LOG(FORM2PRIME) |
| NS Integer | NSINTEGER | LOG(NSINTEGER + 5) |
| Normalized Area | NORMALIZEDAREA | -- |
| Ring Intensity | RINGINTENSITY | -- |
| Ring Thickness | RINGTHICKNESS | -- |

# Statistical Methods

## Modeling Procedures

A number of statistical modeling procedures were implemented during the course of algorithm development. The filtered and not filtered ClearView scan sets were modeled separately, as were the 5 independent diagnosis groups.

These statistical methodologies were each considered as candidates for defining the algorithm. This section describes the 5 statistical methods utilized. While this Section (3.1) describes the basics of each statistical method, Sections 3.2 and 3.3 further describe cross-validation and methods used to evaluate and compare models.

### Logistic Regression

Logistic regression was utilized as a statistical approach to the algorithm modeling. For this method, the dependent variable was the binary representation of subjects’ health status with respect to each of the 5 diagnosis groups. For example, if a subject had a qualifying diagnosis with respect to the cardiovascular system, that subject had a dependent variable of CV=1.

The data for the logistic regression modeling was set-up such that each subject had 78 observations – one observation per sector. Each sector had values for each of the 19 ClearView scan coefficients as the main effects for inclusion in the model as predictors. Additional predictors including age (continuous) and gender were initially considered; however, due of the potential for over-fitting the model to the development dataset, the demographic variables (age, gender) were excluded from the model.

In some instances, preliminary logistic models were executed using main effects and the correlations between some predictors were assessed before implementing the stepwise model selection. With the large number of potential predictors, attention was paid to minimize multicollinearity. If 2 predictors were highly correlated and the model did not benefit from the inclusion of both variables, the stronger predictor was used while the other predictor was not considered for inclusion of the model.

A stepwise model selection procedure was implemented with α = 0.10 significance levels used to determine both entry and exit from the model.

### Naïve Bayes Classifier

Naïve bayes classifiers were also used in the algorithm modeling, as a potential statistical method. While naïve bayes classifiers require an often unrealistic assumption of independence of the predictors, the procedure has been shown to perform quite well even when this assumption is violated. In particular, the naïve bayes classifier can be highly effective for high dimension problems.

For analyses with naïve bayes classifiers, the data was set-up in a wide format. Each subject had 1,482 predictors (19 scan coefficients x 78 sectors) for both the filtered and unfiltered scans. The predictors were set-up to create unique sector-scan coefficient variables. For example, sector1_averageintensity, sector2_averageintensity, etc.

To conduct the naïve bayes regression:

- Fit a univariate general linear regression of Y (dependent variable) on each of the 1,482 predictors (X).  If the regression coefficient is not statistically significantly different from 0, set the regression coefficient to 0.  If the coefficient is statistically significantly different from 0, the coefficient will be used in the next step.

- Create a vector of all resulting univariate regression coefficients.

- Use this coefficient vector and the actual observations for each predictor to get an initial predicted value for each subject.

- Regress the set of Y-values on the set of initial predicted values. The resulting regression coefficient becomes the scaling factor.

- Rescale each of the original regression coefficients from step 2 by multiplying the coefficients by the scaling factor to obtain the final model coefficients.

### Ridge Regression, LASSO, and LARS

In addition to the logistic and naïve bayes regression methods described above, ridge regression, LASSO, and LARS methods were also considered as possible statistical methods for defining the algorithm.

Ridge regression, LASSO, and LARS are all methods for dealing with variable selection in high dimensional situations. While the methodology is different for each method, all 3 sort through a volume of data to determine those predictors that are most informative.

The ridge regression, LASSO, and LARS methods are computationally intense, and were implemented by outside consultants.

In the end, naïve bayes regression was selected over ridge regression, LASSO, LARS, and logistic regression for the primary model. All 5 methods performed well, but the naïve bayes regression was most computationally efficient. In addition to the naïve bayes classification, logistic regression was utilized as a part of the detailed algorithm.

## Cross-Validation

In order to minimize over-fitting, cross-validation was used as part of the modeling process. Cross-validation allowed for an estimation of model performance when used outside of the training dataset.

Two types of cross-validation were implemented. The first method was leave-one-out cross-validation (LOOCV). This method removes a single observation, and then fits the model using all remaining observations. This process is repeated for every observation in the dataset. Essentially, the single observations serve as the validation set at each iteration. The second type of cross-validation used was a generalization of LOOCV, referred to as 6-fold cross-validation (generally, as k-fold cross-validation). For this type of validation, the full training dataset was divided into 6 groups. The model was fit after removing 1 group, using the remaining 5 groups. This process was repeated 5 more times, such that the models were fit 6 times with each group excluded once. The regression coefficients from each of the 6 iterations were averaged to obtain the final model.

The LOOCV method is computationally costly. For this reason, it was not implemented for the naïve bayes modeling. The 6-fold cross-validation method was used for all naïve bayes modeling; LOOCV was used for all other procedures.

## Model Evaluation

The various models were compared using area under the receiver operating characteristic curves (ROC AUC). The ROC curve is the plot of the sensitivity vs. 1-specificity, and the AUC is the probability that the model will rank a randomly chosen case (dependent variable = 1) higher than a randomly chosen control (dependent variable = 0). While there are additional statistics for evaluating predictive models, comparing AUCs is a strong method – particularly when the model does not result in a binary predicted value.

When looking across naïve bayes and logistic regression statistical modeling methodologies, the models with higher AUC values were chosen. The final AUC values of the selected models are presented in Tables 5 and 6 below.

Table 5. Logistic Regression Model Metrics

|  | With Filter AUC | Without Filter AUC |
| --- | --- | --- |
| Cardiovascular | 0.897 | 0.892 |
| GI/Endocrine | 0.823 | 0.828 |
| Renal | 0.896 | 0.899 |
| Hepatic | 0.879 | 0.879 |
| Respiratory | 0.852 | 0.846 |

Table 6. Naïve Bayes Model Metrics

|  | With Filter AUC | Without Filter AUC |
| --- | --- | --- |
| Cardiovascular | 0.838 | 0.792 |
| GI/Endocrine | 0.860 | 0.824 |
| Renal | 0.884 | 0.847 |
| Hepatic | 0.849 | 0.850 |
| Respiratory | 0.828 | 0.800 |

# Calculation of the ClearView Response Scale

The last step in finalizing the ClearView algorithm was to define the Response Scales.

## Epic Score Validator

Before describing the Response Scale, it should be noted that in addition to the statistical methods described within this document, a separate scoring mechanism was developed by Dr. Nancy Rizzo. This mechanism is based upon the manual review of thousands of scans, and utilizes logic that looks for outliers, and then weights and compiles the ClearView scan coefficients into a single value per sector. The single value per sector ranges from 0 to 25.

The details of this mechanism are not described in this document, and will be referred to as the ‘score validator.’

## Sector Scores

### Logistic Regression Scoring

The logistic regression modeling described in Section 3.1.1 above resulted in 5 separate models, 1 for each of the 5 major diagnosis groups. These 5 models were compiled into a single score for each of the 78 sectors, and then scaled to produce a value ranging from 0-25.

To compile the single score per sector, the 5 separate predicted values (for the data with and without the filter) were summed to produce a single score for filtered data, and a single score for the unfiltered data.

Finally, these single scores were transformed to a 0-25 scale by score3.5*35, with any values >25 set to 25.

### Sector Scoring

Based upon Dr. Rizzo’s work with bioelectrophysiology, and ClearView in particular, each of the 78 sectors has been mapped to 1 of 47 organs or systems of the body. This mapping is presented in Appendix B.

After mapping both the logistic regression model results and the score validator results to the 78 sectors to the 47 organs, a comparison was conducted between the 2 scores that were calculated per sector. The greater of the 2 values was selected as the final score for that sector.

Thus, each of the 47 organs had a corresponding 0-25 score. Note that this score is analogous to more traditional GSR values, and represent the detailed components of the primary model results.

## Primary Scores

Naïve bayes models results were used to generate the 5 primary Response Scale results (for scans with and without the filter).

To transform the 5 naïve bayes model results to a Response Scale ranging from 0 to 5, the following calculations were implemented for each of the models.  First, the coefficients from the naïve bayes models for each diagnosis group were used to calculate a new score (newdx) for the diagnosis group of interest. The coefficients are in Table 7.  Next, the score was scaled to 0-5 for the final Response Scale.

Table 7: Conversion of Naïve Bayes Model Output to Response Scale

|  |  |  |  | Original Naïve Bayes Model Output Per Diagnosis Group |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Filtered | Dependent | Intercept | CV |  | GI | Hepatic | Renal | Respiratory |
| 0 | cv | -6.91 | -3.43 |  | 0.00 | -2.03 | 11.54 | 0.00 |
| 0 | gi | -10.42 | -29.41 |  | 33.35 | 8.97 | -22.88 | 16.08 |
| 0 | kid | -12.54 | -38.64 |  | 30.71 | 12.39 | -26.59 | 34.98 |
| 0 | liv | -4.42 | 0.00 |  | -7.85 | -3.30 | 12.63 | 0.00 |
| 0 | resp | -3.72 | 2.58 |  | -13.53 | -4.05 | 13.97 | 5.79 |
| 1 | cv | 5.88 | -5.07 |  | 2.03 | -5.06 | 10.30 | 3.17 |
| 1 | gi | -4.83 | -11.24 |  | 13.38 | -1.07 | -1.37 | 1.79 |
| 1 | kid | 2.18 | -9.60 |  | 2.54 | -4.64 | 8.93 | 7.95 |
| 1 | liv | -0.82 | -14.73 |  | 13.95 | 11.87 | -3.55 | 0.00 |
| 1 | resp | 1.82 | -7.25 |  | 2.19 | -6.90 | 7.55 | 7.59 |

As an example, for the Cardiovascular Response Scale for the filtered data:

newcv = 5.88 + (nbcv)* (-5.07) + (nbgi)*( 2.03) + (nbliv)* (-5.06) + (nbkid)*(10.30)+(nbresp)*(3.17),

where newcv is an intermediary score, and nbcv, nbgi, nbliv, nbkid, and nbresp are the initial results from the naïve bayes models.

Then, after implementing the calculations and obtaining the “newdx” values (the score for each group after modeling), convert to 0-5 by:

30*enewdx1+enewdx*15.

This value represents the final Response Scale value.

# Validation of Implementation

SAS v9.2 was used for all statistical modeling (with the exception of the creation of the score validator) (SAS Institute Inc., Cary, NC, USA). The final algorithm was then implemented in the ClearView system.

To ensure that the SAS results and ClearView output match, a point-by-point comparison was conducted. Fifteen subjects from the development cohort, 3 subjects from each of the 5 studies, were randomly selected for the comparison. The 15 subjects selected for this comparison are listed in Appendix C.

Output from the ClearView system was extracted as CSV files, and imported to SAS. The SAS and ClearView results were merged by subject id, and the differences were calculated for each score produced – the 47 organs, as well as the 5 primary Response Scale values. No differences were found for any of the scores, verifying that the algorithm developed in SAS was correctly transferred to the ClearView software.

# Appendix A: Description of ClearView Scan Coefficients

## Sector Area and Normalized Area

Sector Area parameter of a particular radial sector can be computed, such as for one or more radial sectors of the energized image.  The Sector Area of a particular sector can be determined as the number of pixels within a particular sector and within the fitted ellipse, having an intensity exceeding a specified value, such as exceeding a specified value of the Background Noise Level. Normalized area uses the calculated sector area and corrects it for sector size.

## Average Intensity

The Average Intensity of a particular sector can be determined by dividing the sum of intensities of all pixels in a particular sector by the number of pixels given by the Sector Area for that sector.

## Entropy

An Entropy for a particular sector can be computed, such as by first computing an Entropy for each individual profile within that particular sector, and then averaging or otherwise determining a central tendency of each individual profile to obtain a composite profile for that particular sector.  For various pixel positions i along the radial profile (where the integer i = 1, 2, . . n, and n is the total number of pixels in the radial profile), the Entropy can be expressed as a radial vector E given by the following relationship:

E = - ∑iµilog10µi,

where µi = ΔIi/∑i ΔIi, and

ΔIi is pixel intensity above the Background Noise Level.

## Form

A Form parameter of a particular sector can be computed, such as for one or more radial sectors of the energized image.  Form = 4πL/S(1), calculated for concentric region having respective areas A1, A2, and A3 where S(1) =  ΔIi/∑i ΔImax,

where ΔIi is the pixel intensity above the Background Noise Level,

and L is the length as given in pixels of intensity described by S(1).

## Form 2 and Form2’

Form 2 can be calculated using a similar calculation as Form; however it can be carried out for only the concentric region that is greater than that associated with A3, namely the concentric region associated with A4 and greater radial distances. Where Form 2 uses the strict sector area and Form 2’ uses the normalized area.

## Fractal

Fractal dimension parameter of a particular sector can be computed, such as for one or more radial sectors of the energized image.  The Fractal parameter can be determined by computing a mathematical fractal dimension, such as using a box-counting method for a two-dimensional area.  The Fractal parameter can be represented by:

Fractal = 2πL/Rav,

where 2πL is perimeter length (in pixel count) of the radial image sector, and

where Rav=(1/n)*∑inRi, and Ri is the Inner Radius.

## NS Parameter

A NS parameter of a particular sector can be computed, such as for one or more radial sectors of the energized image.  NS can provide a comparison measure between a subject’s image sector and a corresponding sector of a subject-specific calibration image (e.g., a calibration image that has been taken on the same day as the subject images).  The NS comparison can be determined both with and without the dielectric 103 in place.  In an example, the NS parameter can be determined for a particular sector using the following relationship:

NS = [(Iav (subject image sector) ÷ Iav(calibration image sector)) - 0.5]  ÷ [(log (number of active pixels of subject image sector) ÷ (number of total pixels of subject image sector + epsilon)) ÷ (log (number of active pixels of calibration image sector) ÷ (number of total pixels of calibration image sector + epsilon))],

where Iav = average intensity of the pixels in the radial sector of the image, and epsilon can be set to a value (e.g., epsilon = 10-4) to help ensure stability, and the value 0.5 can be subtracted for normalization.

# Appendix B: Mapping 78 Sectors to 47 Organs

| Order on Report | Organ | Sector | Side | System Label |
| --- | --- | --- | --- | --- |
| 2 | Eye (R) | 1L-1 | L | Sensory & Skeletal Systems |
| 4 | Ear/Nose/Sinus (R) | 1L-2 | L | Sensory & Skeletal Systems |
| 6 | Jaw/Teeth (R) | 1L-3 | L | Sensory & Skeletal Systems |
| 22 | Throat/Larynx/Trachea/ | 1L-4 | L | Cardiovascular System |
| 5 | Jaw/Teeth (L) | 1L-5 | L | Sensory & Skeletal Systems |
| 3 | Ear/Nose/Sinus (L) | 1L-6 | L | Sensory & Skeletal Systems |
| 1 | Eye (L) | 1L-7 | L | Sensory & Skeletal Systems |
| 16 | Cerebral Cortex | 1L-8 | L | Nervous & Immune Systems |
| 2 | Eye (R) | 1R-1 | R | Sensory & Skeletal Systems |
| 4 | Ear/Nose/Sinus (R) | 1R-2 | R | Sensory & Skeletal Systems |
| 6 | Jaw/Teeth (R) | 1R-3 | R | Sensory & Skeletal Systems |
| 22 | Throat/Larynx/Trachea/ | 1R-4 | R | Cardiovascular System |
| 5 | Jaw/Teeth (L) | 1R-5 | R | Sensory & Skeletal Systems |
| 3 | Ear/Nose/Sinus (L) | 1R-6 | R | Sensory & Skeletal Systems |
| 1 | Eye (L) | 1R-7 | R | Sensory & Skeletal Systems |
| 16 | Cerebral Cortex | 1R-8 | R | Nervous & Immune Systems |
| 40 | Descending Colon | 2L-1 | L | Gastro Intestinal & Large |
| 41 | Sigmoid Colon | 2L-2 | L | Gastro Intestinal & Large |
| 42 | Rectum | 2L-3 | L | Gastro Intestinal & Large |
| 11 | Coccyx/Pelvis | 2L-4 | L | Sensory & Skeletal Systems |
| 10 | Sacrum | 2L-5 | L | Sensory & Skeletal Systems |
| 9 | Lumbar Spine | 2L-6 | L | Sensory & Skeletal Systems |
| 8 | Thoracic Spine | 2L-7 | L | Sensory & Skeletal Systems |
| 7 | Cervical Spine | 2L-8 | L | Sensory & Skeletal Systems |
| 39 | Transverse Colon | 2L-9 | L | Gastro Intestinal & Large |
| 7 | Cervical Spine | 2R-1 | R | Sensory & Skeletal Systems |
| 8 | Thoracic Spine | 2R-2 | R | Sensory & Skeletal Systems |
| 9 | Lumbar Spine | 2R-3 | R | Sensory & Skeletal Systems |
| 10 | Sacrum | 2R-4 | R | Sensory & Skeletal Systems |
| 11 | Coccyx/Pelvis | 2R-5 | R | Sensory & Skeletal Systems |
| 31 | Blind Gut | 2R-6 | R | Stomach & Small Intestine |
| 37 | Appendix | 2R-7 | R | Gastro Intestinal & Large |
| 38 | Ascending Colon | 2R-8 | R | Gastro Intestinal & Large |
| 39 | Transverse Colon | 2R-9 | R | Gastro Intestinal & Large |
| 27 | Cardiovascular Circulation | 3L-1 | L | Cardiovascular System |
| 44 | Kidney - Left | 3L-2 | L | Renal & Reproductive |
| 35 | Liver | 3L-3 | L | Endocrine & Metabolism |
| 32 | Abdominal Region | 3L-4 | L | Stomach & Small Intestine |
| 18 | Immune System | 3L-5 | L | Nervous & Immune Systems |
| 21 | Thorax Respiratory | 3L-6 | L | Respiratory System |
| 17 | Cerebral Vessels | 3L-7 | L | Nervous & Immune Systems |
| 21 | Thorax Respiratory | 3R-1 | R | Respiratory System |
| 18 | Immune System | 3R-2 | R | Nervous & Immune Systems |
| 36 | Gallbladder | 3R-3 | R | Endocrine & Metabolism |
| 35 | Liver | 3R-4 | R | Endocrine & Metabolism |
| 45 | Kidney - Right | 3R-5 | R | Renal & Reproductive |
| 27 | Cardiovascular Circula | 3R-6 | R | Cardiovascular System |
| 17 | Cerebral Vessels | 3R-7 | R | Nervous & Immune Systems |
| 13 | Hypothalamus | 4L-1 | L | Nervous & Immune Systems |
| 12 | Nervous System | 4L-2 | L | Nervous & Immune Systems |
| 19 | Spleen | 4L-3 | L | Nervous & Immune Systems |
| 47 | Genitourinary System | 4L-4 | L | Renal & Reproductive |
| 43 | Adrenal | 4L-5 | L | Renal & Reproductive |
| 34 | Pancreas | 4L-6 | L | Endocrine & Metabolism |
| 33 | Thyroid | 4L-7 | L | Endocrine & Metabolism |
| 14 | Pituitary | 4L-8 | L | Nervous & Immune Systems |
| 15 | Pineal | 4L-9 | L | Nervous & Immune Systems |
| 14 | Pituitary | 4R-1 | R | Nervous & Immune Systems |
| 33 | Thyroid | 4R-2 | R | Endocrine & Metabolism |
| 34 | Pancreas | 4R-3 | R | Endocrine & Metabolism |
| 43 | Adrenal | 4R-4 | R | Renal & Reproductive |
| 47 | Genitourinary System | 4R-5 | R | Renal & Reproductive |
| 19 | Spleen | 4R-6 | R | Nervous & Immune Systems |
| 12 | Nervous System | 4R-7 | R | Nervous & Immune Systems |
| 13 | Hypothalamus | 4R-8 | R | Nervous & Immune Systems |
| 15 | Pineal | 4R-9 | R | Nervous & Immune Systems |
| 25 | Heart (Left Side) | 5L-1 | L | Cardiovascular System |
| 46 | Urokidney | 5L-2 | L | Renal & Reproductive |
| 20 | Respiratory/Mammary | 5L-3 | L | Respiratory System |
| 29 | Jejunum | 5L-4 | L | Stomach & Small Intestine |
| 26 | Heart (Right Side) | 5L-5 | L | Cardiovascular System |
| 23 | Coronary Vessels | 5L-6 | L | Cardiovascular System |
| 28 | Duodenum | 5R-1 | R | Stomach & Small Intestine |
| 30 | Ileum | 5R-2 | R | Stomach & Small Intestine |
| 20 | Respiratory/Mammary | 5R-3 | R | Respiratory System |
| 46 | Urokidney | 5R-4 | R | Renal & Reproductive |
| 24 | Heart | 5R-5 | R | Cardiovascular System |
| 23 | Coronary Vessels | 5R-6 | R | Cardiovascular System |

# Appendix C: Subjects Selected for SAS vs ClearView Verification

| Subject | Patient Treatment ID | Age | Gender |
| --- | --- | --- | --- |
| DV001-0006024 | 2139 | 52 | 1 |
| DV001-0006435 | 2141 | 62 | 0 |
| DV001-0009611 | 2142 | 28 | 0 |
| EHI-0001967 | 2143 | 59 | 0 |
| EHI-0005063 | 2147 | 47 | 0 |
| EHI-0005800 | 2148 | 53 | 0 |
| EPIC-5065014 | 2150 | 21 | 0 |
| EPIC-9563536 | 2152 | 57 | 0 |
| EPIC-9757383 | 2153 | 61 | 1 |
| HVS-0002406 | 2154 | 79 | 1 |
| HVS-0006235 | 2156 | 90 | 0 |
| HVS-0008555 | 2157 | 55 | 1 |
| RPT-0006596 | 2158 | 44 | 1 |
| RPT-0008692 | 2247 | 64 | 1 |
| RPT-0008706 | 2248 | 45 | 1 |