Originator:

| Name | Job Title | Signature | Date |
| --- | --- | --- | --- |
| ##### Andrew Mason | Chief Technology Officer |  |  |

Approvals:

| Name | Job Title | Signature | Date |
| --- | --- | --- | --- |
| ##### Brian Tompson | Software Engineer |  |  |
| ##### Don Pegg | ##### Senior Electrical Engineer |  |  |
| ##### Andrea Miller | ##### Quality Assurance Manager |  |  |

- Purpose

This report summarizes the testing performed to validate the application level firmware interface.

- Executive Summary

The application level firmware interface is the link between the ClearView software application and the ClearView device. This interface is responsible for sending commands to the firmware through a USB communication channel. These commands instruct the firmware to take specific actions on behalf of the ClearView software. This protocol is intended to exercise this interface to verify that the interaction between the ClearView software and the ClearView device act as designed.

- Protocol Execution and Results

All validation activities were conducted by EPIC staff on 5/18/12.  The technician completed the protocol, recording the results directly on the protocol.  The original protocol is located in Attachment A.  A CD of electronic files generated during the validation testing is located in Attachment B.  All results met the expected outcome.

  - Two minor deviations in the protocol instructions were recorded directly onto the protocol.  These changes were to ensure that the protocol provided proper instructions to the technician executing the protocol.  No other deviations were identified during the execution of this protocol.  These changes will be made to the protocol to ensure that future execution of this verification testing is performed accurately.

  - No non-conformances were encountered.

- Conclusions

The application level firmware interface in version R04.00.00 is considered validated for use in IDE investigations.

- Attachments

Attachment A- Original Signed Protocol

Attachment B- CD of electronic files