Approval Signatures:

Author:

| Department | Name | Signature | Date |
| --- | --- | --- | --- |
| Engineering | Brian Tompson |  |  |

Approval:

| Engineering | Don Pegg |  |  |
| --- | --- | --- | --- |
| IS | Andrew Mason |  |  |
| QA | Andrea Miller |  |  |

# PWM Problem Report

# for the

# EPIC ClearView Device

# 03-February-2012

# Purpose

This document describes the resolution to an issue with the ClearView device PWM signal that caused excessive Boost Voltages to be produced inside the unit.

# Scope

This document describes the discovery of the excessive Boost Voltage during testing, the diagnosis of the root cause, and the resolution to the problem.  Testing results are included that verify that the excessive Boost Voltage is no longer present.

# Overview

During testing of firmware version R03.99.19 on the ClearView device an audible ‘pop’ was sometimes heard coming from the device.

A recurring overvoltage condition of the secondary Boost Voltage was discovered.  The overvoltage conditions were caused by an improper PWM0 signal generated by the MCU.  The PWM0 signal controls the Boost Voltage inside the ClearView device.

Firmware version R03.99.29 was created to better control the PWM0 signal.  Testing of the R03.99.29 confirms that the PWM0 no longer produces the improper signal, and excessive secondary Boost Voltage is no longer generated.

# Detailed Problem Description

## Excessive Boost Voltage Discovery

Many times while testing the ClearView device an audible ‘pop’ could be heard coming from the device when an exposure was taken.  When some of these ‘pop’ events occurred images were captured.  These images showed a larger than normal energy discharge.  See Figure 4-1, and Figure 4-2.

The ClearView firmware version R03.99.19 was initially used during this testing.

Figure 4-1 Pop Event - ClearView Camera

Figure 4-2 Pop Event - External Camera

## Problem Analysis

The ClearView device uses a two stage Boost Voltage to energize a glass plate during image capture.  The amount of Boost Voltage is controlled by a Pulse Width Modulated (PWM) signal generated by the MCU inside the device.  This signal is referred to as ‘PWM0’.

Additional testing determined that prior to a ‘pop’ event the duty cycle pulse of the PWM0 signal was excessively long.  The correct duty cycle is 14 microseconds.  Duty cycles as long as 3.84 milliseconds were measured prior to a ‘pop’ event.  These long duty cycle events caused excessive secondary Boost Voltages to be generated.  When the excessive Boost Voltage was discharged during the following exposure, the ‘pop’ event would occur.

Initially it was thought that the ‘pop’ event only occurred during the first exposure after the ClearView device was initially powered up.  Firmware version R03.99.21 was created to eliminate long PWM0 pulses during the first exposure after power up.  This version of firmware used the MCU’s PWM hardware during initialization to generate a PWM signal inside the MCU.  This internal PWM signal never left the MCU, therefore the Boost Voltage was not affected.  After initialization, the PWM0 signal was controlled as in the R03.99.19 version of firmware.  This did eliminate the ‘pop’ event during the first exposure after power up.

As testing continued, it was determined that the ‘pop’ event additionally occurred randomly after a number of exposures had been taken.  The ‘pop’ event, now known to be a secondary Boost Voltage overvoltage event, typically happened within 20 minutes of repeated exposure testing.

Analysis of the firmware design determined that it at the end of an exposure it was possible for the internal PWM timer to be turned off during the active duty cycle portion of the PWM0 signal.  The PWM0 signal was subsequently commanded off by the firmware, but the timing of when this occurred was not assured.

The PWM0 signal needs to be better controlled to prevent any long pulses from occurring.  During the analysis it was also determined that a shorter than normal PWM0 pulse would have no undesired affect on the Boost Voltage.

## Problem Resolution

Instead of turning off the PWM timer at the end of an exposure, the PWM0 signal will be gated by the firmware.  The MCU’s PWM hardware would continuously generate the PWM signal internally, but the signal would only reach the PWM0 output pin during an exposure.  By not turning off the PWM timer, the duty cycle of the PWM signal is always controlled by the MCU hardware.  When the firmware ends an exposure it can no longer cause a long PWM0 pulse.

By gating the PWM0 output, the duty cycle of the pulse could now be shorter than normal if the firmware ends an exposure during the active duty cycle portion of the PWM0 signal. Since only long PWM0 pulses cause the excessive secondary Boost Voltages to be generated, a short pulse would not cause a Boost Voltage overvoltage event.

Firmware version R03.99.29 was created to operate the MCU’s PWM generator continuously and gate the PWM0 signal.

## Verification

A ClearView device was instrumented to detect a PWM0 duty cycle in excess of 20 microseconds, or excessive secondary Boost Voltages.  Firmware version R03.99.29 installed.

The ClearView device was tested for an entire day using the default settings of; 1100 Hz PWM frequency, 0.5 second exposure, and 100% Boost Voltage.  No secondary Boost Voltage overvoltage events, or long PWM0 pulse events were detected.

Next, every selectable PWM frequency was tested for at least an hour, using a 0.5 second exposure and 100% Boost Voltage.  Again, no secondary Boost Voltage overvoltage events, or long PWM0 events were detected.

Every selectable exposure duration was tested using a PWM frequency of 1100 Hz and 100% Boost Voltage setting.  No secondary Boost Voltage overvoltage events or long PWM0 events were detected.

Testing was performed to detect a Boost Voltage overvoltage event, or long PWM0 pulse event, during the first exposure after initially powering up the ClearView device.  No secondary Boost Voltage overvoltage events, or long PWM0 events were detected.

During the problem analysis testing it was also noticed that the MCU inside the ClearView device would reset itself when a ‘pop’ event occurred.  No MCU resets were detected during testing.

Document Revision History

| Version Number: | Description of Change: | Date: | Updated by: |
| --- | --- | --- | --- |
| 000 | Initial document creation | 5/21/12 | B. Tompson |