Approval Signatures:

| Department | Name | Signature | Date |
| --- | --- | --- | --- |
| Author | Dan Miller |  |  |

Approval:

| OPS | Mike Stowell |  |  |
| --- | --- | --- | --- |
| Exec | Nancy Rizzo |  |  |
| IS | Andrew Mason |  |  |
| QA | Scott Pletzer |  |  |

- Purpose

This report summarizes the testing performed to verify the effect of changing the calibration acceptance limits within the ClearView Software.

- Executive Summary

The verification activities demonstrate that the changes to the ClearView Software to increase the limits on the outer ring from +/- 3 standard deviations to +3/-4 standard deviations from the expected pixel intensities.  This study loaded known calibration images which lie within the new acceptance criteria (high, middle and low within the acceptable pixel intensity range) and subsequently loaded known human subject images.  A comparison was made between the Naïve Bayes scores for all subjects across the different calibration image ranges.  The comparison demonstrated that the human subject Naïve Bayes (NB) values did not vary by more than 0.3 of the total scale of 0 to 5.  Therefore, the change to the calibration acceptance limit is considered verified for use.

- Protocol Execution and Results

All validation activities were conducted by EPIC staff on 7/11/12.  The technician completed the protocol, recording the results directly on the protocol.  The original protocol is located in Attachment A.  A CD of electronic files generated during the validation testing is located in Attachment B.  All ClearView printouts generated during the execution of the protocol are located in Attachment C.  The human subject names were used when conducting this testing.  As such, the names have been redacted from the original printouts (removed by covering with black marker) and the subject number recorded on the printouts.  Additionally, the subject files were modified in the ClearView database to reflect the subject name (e.g., First Name ENG-003, Last Name Subject 1).

  - Deviations - No deviations were identified during the execution of this protocol.

  - Non-conformances – All protocol requirements were met and there were no non-conformances identified.

The test required that three different sets of calibration images were used to evaluate the effect of the proposed change to calibration acceptance criteria.  The pixel intensities of these three sets were established by selecting brightness and gain camera settings that produce results 1) near the lower limit of the new calibration acceptance criteria, 2) near the upper limit of the new calibration acceptance criteria, and 3) results from calibration at the optimum settings for human images.  A copy of these results is located in Attachment D.

The results of the comparison are summarized in Table 1 below.  A comparison was made between the NB values obtained using the mid-range calibration images and the NB values obtained when using the low and high-range calibration images.  Changes are identified by shading the cell grey.

Table 1

Subject 1

| Settings | B85G58 | B85G55 | B100G63 | Variation |  | B85G58 | B85G55 | B100G63 | Variation |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Organ System | Physical |  |  |  |  | Autonomic |  |  |  |
| Cardiovascular | 3.2 | 3.2 | 3.1 | -0.1 |  | 2.1 | 2.1 | 1.9 | -0.2 |
| Respiratory | 2.6 | 2.6 | 2.6 | 0.0 |  | 1.4 | 1.4 | 1.3 | -0.1 |
| Hepatic | 0.2 | 0.2 | 0.3 | +0.1 |  | 0.9 | 0.9 | 0.9 | 0.0 |
| Gastrointestinal | 2.5 | 2.5 | 2.5 | 0.0 |  | 1.8 | 1.9 | 1.6 | -0.1/-0.3 |
| Renal | 1.3 | 1.3 | 1.3 | 0.0 |  | 0.4 | 0.4 | 0.3 | -0.1 |
| Range of Variation +/- |  |  |  | 0.0 – 0.1 |  | Range of Variation +/- |  |  | 0.0 – 0.3 |

Subject 2

| Settings | B85G58 | B85G55 | B100G63 | Variation |  | B85G58 | B85G55 | B100G63 | Variation |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Organ System | Physical |  |  |  |  | Autonomic |  |  |  |
| Cardiovascular | 2.7 | 2.7 | 2.6 | +0.1 |  | 2.1 | 2.2 | 1.9 | -0.1/-0.3 |
| Respiratory | 1.9 | 1.9 | 1.9 | 0.0 |  | 1.4 | 1.4 | 1.3 | -0.1 |
| Hepatic | 0.5 | 0.4 | 0.5 | +0.1 |  | 0.9 | 0.9 | 0.9 | 0.0 |
| Gastrointestinal | 2.6 | 2.6 | 2.6 | 0.0 |  | 1.3 | 1.3 | 1.1 | -0.2 |
| Renal | 0.8 | 0.8 | 0.9 | +0.1 |  | 0.1 | 0.2 | 0.1 | -0.1 |
| Range of Variation +/- |  |  |  | 0.0 – 0.1 |  | Range of Variation +/- |  |  | 0.0 – 0.3 |

Subject 3

| Settings | B85G58 | B85G55 | B100G63 | Variation |  | B85G58 | B85G55 | B100G63 | Variation |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Organ System | Physical |  |  |  |  | Autonomic |  |  |  |
| Cardiovascular | 2.0 | 2.0 | 2.0 | 0.0 |  | 1.6 | 1.6 | 1.5 | -0.1 |
| Respiratory | 1.4 | 1.4 | 1.4 | 0.0 |  | 1.2 | 1.2 | 1.2 | 0.0 |
| Hepatic | 0.3 | 0.3 | 0.3 | 0.0 |  | 0.8 | 0.8 | 0.8 | 0.0 |
| Gastrointestinal | 2.0 | 2.0 | 2.0 | 0.0 |  | 1.5 | 1.6 | 1.4 | -0.1/-0.2 |
| Renal | 0.6 | 0.6 | 0.6 | 0.0 |  | 0.3 | 0.4 | 0.3 | -0.1 |
| Range of Variation +/- |  |  |  | None |  | Range of Variation +/- |  |  | 0.0 – 0.2 |

Subject 4

| Settings | B85G58 | B85G55 | B100G63 | Variation |  | B85G58 | B85G55 | B100G63 | Variation |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Organ System | Physical |  |  |  |  | Autonomic |  |  |  |
| Cardiovascular | 1.0 | 1.0 | 0.9 | -0.1 |  | 2.1 | 2.1 | 1.9 | -0.2 |
| Respiratory | 1.3 | 1.3 | 1.3 | 0.0 |  | 1.5 | 1.6 | 1.5 | -0.1 |
| Hepatic | 0.3 | 0.2 | 0.3 | +0.1 |  | 1.4 | 1.4 | 1.4 | 0.0 |
| Gastrointestinal | 2.5 | 2.5 | 2.5 | 0.0 |  | 1.4 | 1.5 | 1.2 | -0.1/-0.3 |
| Renal | 0.5 | 0.5 | 0.5 | 0.0 |  | 0.2 | 0.2 | 0.1 | -0.1 |
| Range of Variation +/- |  |  |  | 0.0 – 0.1 |  | Range of Variation +/- |  |  | 0.0 – 0.3 |

Subject 5

| Settings | B85G58 | B85G55 | B100G63 | Variation |  | B85G58 | B85G55 | B100G63 | Variation |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Organ System | Physical |  |  |  |  | Autonomic |  |  |  |
| Cardiovascular | 2.0 | 2.0 | 1.9 | -0.1 |  | 1.5 | 1.6 | 1.5 | -0.1 |
| Respiratory | 1.7 | 1.7 | 1.7 | 0.0 |  | 1.2 | 1.2 | 1.1 | -0.1 |
| Hepatic | 0.2 | 0.2 | 0.2 | 0.0 |  | 0.8 | 0.8 | 0.8 | 0.0 |
| Gastrointestinal | 2.0 | 2.0 | 2.0 | 0.0 |  | 1.6 | 1.7 | 1.5 | -0.1/-0.2 |
| Renal | 0.7 | 0.7 | 0.7 | 0.0 |  | 0.3 | 0.3 | 0.2 | -0.1 |
| Range of Variation +/- |  |  |  | 0.0 – 0.1 |  | Range of Variation +/- |  |  | 0.0 – 0.2 |

The human subject physical NB values had less variation than the autonomic NB values.  Thirty-seven (37) of the fifty (50) or 84% of the physical NB values had no variation in when using calibration images at various points along the proposed acceptance range.  Eight (8) of the fifty (50) or 16% of the physical NB values had only a 0.1 point variation.  Additionally, variation was seen in only two of the NB values obtained when using the low-range calibration images.  The remaining instances of variation were seen in the NB values obtained when using the high-range calibration images.

The variation was slightly higher for the human autonomic NB values.   The average variation between all three sets is still only 0.1.  However, the range of variation increased to include - 0.3.  All changes in the autonomic NB values were to lower the NB value; no increases in NB values were identified.   Twenty-two (22) of the fifty (50) or 44% of the autonomic NB values had no variation in when using calibration images at various points along the proposed acceptance range.  Twenty-eight (28) of the fifty (50) or 56% of the autonomic NB values had variation.  Additionally, variation was seen more frequently in the NB values obtained when using the high-range calibration images (a total of nineteen, 19, of the values that had variation).  The remaining instances of variation were seen in the NB values obtained when using the low-range calibration images, a total of 9 values.

Table 2 below summarizes the variation from the values obtained during the execution of this protocol to the original values obtained using these images during the SN0986 performance qualification ENG-017.002.  Changes are identified by shading the cell grey.

Table 2

Subject 1

| Settings | B85G58 | B85G55 | B100G63 |  | B85G58 | B85G55 | B100G63 |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Organ System | Physical |  |  |  | Autonomic |  |  |
| Cardiovascular | 0.0 | 0.0 | -0.1 |  | +0.1 | +0.1 | -0.1 |
| Respiratory | 0.0 | 0.0 | 0.0 |  | 0.0 | 0.0 | -0.1 |
| Hepatic | 0.0 | 0.0 | -0.1 |  | 0.0 | 0.0 | 0.0 |
| Gastrointestinal | 0.0 | 0.0 | 0.0 |  | +0.1 | +0.2 | -0.1 |
| Renal | 0.0 | 0.0 | 0.0 |  | +0.1 | +0.1 | 0.0 |

Subject 2

| Settings | B85G58 | B85G55 | B100G63 |  | B85G58 | B85G55 | B100G63 |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Organ System | Physical |  |  |  | Autonomic |  |  |
| Cardiovascular | +0.1 | +0.1 | 0.0 |  | +0.2 | +0.3 | 0.0 |
| Respiratory | 0.0 | 0.0 | 0.0 |  | +0.1 | +0.1 | 0.0 |
| Hepatic | 0.0 | -0.1 | 0.0 |  | 0.0 | 0.0 | 0.0 |
| Gastrointestinal | 0.0 | 0.0 | 0.0 |  | +0.2 | +0.2 | 0.0 |
| Renal | -0.1 | -0.1 | 0.0 |  | 0.0 | +0.1 | 0.0 |

Subject 3

| Settings | B85G58 | B85G55 | B100G63 |  | B85G58 | B85G55 | B100G63 |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Organ System | Physical |  |  |  | Autonomic |  |  |
| Cardiovascular | 0.0 | 0.0 | 0.0 |  | +0.1 | +0.1 | 0.0 |
| Respiratory | 0.0 | 0.0 | 0.0 |  | 0.0 | 0.0 | 0.0 |
| Hepatic | 0.0 | 0.0 | 0.0 |  | 0.0 | 0.0 | 0.0 |
| Gastrointestinal | 0.0 | 0.0 | 0.0 |  | +0.1 | +0.2 | 0.0 |
| Renal | 0.0 | 0.0 | 0.0 |  | 0.0 | +0.1 | 0.0 |

Subject 4

| Settings | B85G58 | B85G55 | B100G63 |  | B85G58 | B85G55 | B100G63 |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Organ System | Physical |  |  |  | Autonomic |  |  |
| Cardiovascular | +0.1 | +0.1 | 0.0 |  | +0.2 | +0.2 | 0.0 |
| Respiratory | 0.0 | 0.0 | 0.0 |  | 0.0 | +0.1 | 0.0 |
| Hepatic | 0.0 | -0.1 | 0.0 |  | 0.0 | 0.0 | 0.0 |
| Gastrointestinal | 0.0 | 0.0 | 0.0 |  | +0.1 | +0.2 | -0.1 |
| Renal | 0.0 | 0.0 | 0.0 |  | +0.1 | +0.1 | 0.0 |

Subject 5

| Settings | B85G58 | B85G55 | B100G63 |  | B85G58 | B85G55 | B100G63 |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Organ System | Physical |  |  |  | Autonomic |  |  |
| Cardiovascular | +0.1 | +0.1 | 0.0 |  | 0.0 | +0.1 | 0.0 |
| Respiratory | -0.1 | -0.1 | -0.1 |  | +0.1 | +0.1 | 0.0 |
| Hepatic | 0.0 | 0.0 | 0.0 |  | 0.0 | 0.0 | 0.0 |
| Gastrointestinal | 0.0 | 0.0 | 0.0 |  | +0.1 | +0.2 | 0.0 |
| Renal | 0.0 | 0.0 | 0.0 |  | 0.0 | 0.0 | -0.1 |

- Evaluation

This study was conducted to establish the effect of the proposed change to the calibration acceptance criteria.  Although there was variation in the NB values, the variation was not more than 0.3 of the total scale of 0 to 5.  A clinical study, DV-002, was conducted to evaluate the repeatability of the ClearView measurements.  The data collected determined the largest interval limit for change was <0.45.  This amount of variation in reported values was determined to be within a range that is not considered clinically meaningful (Reference Summary of DV-002 Statistical Report ClearView Repeat Long-term Profile Testing).  Therefore, the variation identified in this study (i.e., ≤0.3) falls within the range that is not considered clinically meaningful.

All variation was found to be in the range of acceptance established in the protocol (no more than 0.5 point change in NB values using different calibration images).    The majority of the variation (i.e., only 27 of the 36 NB values which showed variation or 75%) was observed when using the calibration images from the high side of the range.  Also, the variation observed when using calibration images from the high side of the acceptance criteria creates greater variation (i.e., up to 0.3 points) than that observed when using calibration images near the low end of the acceptance criteria (0.1 point).  Given that the high side of the calibration acceptance criteria is not being changed, the variation in NB values when accepting calibration images near the newly proposed low end of the acceptance criteria demonstrates less variation than what is currently acceptable.  Therefore, the proposed change in calibration acceptance criteria has been demonstrated as having no adverse effect on the final human subject NB values and can be considered acceptable for implementation.

- Conclusions

The proposed change to the calibration acceptance criteria has been established to not significantly affect human subject results.  Therefore, it is recommended to implement the proposed change within the ClearView Software.

- Attachments

Attachment A- Original Signed Protocol

Attachment B- CD of electronic files

Attachment C- ClearView Report Printouts

Attachment D- Calibration Image Intensity Values