Approval Signatures:

Author:

| Department | Name | Signature | Date |
| --- | --- | --- | --- |
| Author - ENG | Dan Miller |  |  |

Approval:

| OPS | Mike Stowell |  |  |
| --- | --- | --- | --- |
| Exec | Nancy Rizzo |  |  |
| QA | Scott Pletzer |  |  |

- Purpose

This report summarizes the testing performed to investigate the glass failure originally noted on PADT’s Nonconformances Report (NCR) #07703.  This report does not finalize the investigation, but is intended to document the investigation to this point.

- Executive Summary

After a catastrophic failure in the glass electrode of a ClearView Device assembled for engineering testing, this investigation was opened to identify the appropriate cause and implement corrective and preventive action.  This report summarizes the action taken to identify glass that will not fail catastrophically.  Two different test methods were developed to identify glass assemblies which will not fail catastrophically.  Four ClearView devices assembled with passing glass assemblies have been tested repeatedly without failure.  Until the root cause of the failures can be identified, this report will be used to implement testing of all glass assemblies prior to manufacture in devices.

- Background

The ClearView Device is the hardware component of the ClearView System.  The device is controlled by the user through a software interface, the ClearView Software.  The system measures the electrical resistance of a subject’s finger in contact with a glass electrode.  A series of painless electrical impulses are applied to the glass electrode generating a localized electromagnetic field around the finger.  The glass electrode (part # 06617) consists of a plastic housing, simple electrode board and the glass assembly (part # 1282-1016).  The glass assembly consists of two round discs of glass where one disc has one side coated with a conductive indium tin oxide (ITO) layer.  The two discs are laminated together to form the final assembly.  This study documents the investigation of a catastrophic failure of the glass assembly in the engineering unit built by EPIC’s contract manufacturer.

The failure appears to be the result of fusing of the conductive ITO layer upon the application of energy during the initial exposures of the electrical impulses.  The fusing is seen as fine lines made visible due to the melting of the lamination material against the ITO surface.  Figure 1 is a representative picture of the catastrophic failure.

Figure 1

- Original NCR Testing

An NCR was originated by EPIC’s contract engineering firm, PADT, Inc., to address this failure.  A copy of NCR # 07703 is located in Attachment A of this report.  The NCR provided MRB instructions for testing the glass prior to placement into the electrode assembly.  The test method established on the NCR is as follows:

  - A test fixture (see Attachment B) was built to apply energy to the glass assembly before final assembly into the electrode housing.  The glass assembly is placed into the test fixture.

  - The boost voltage is set to 100% (approximately 129 volts) using the ClearView software and a duty cycle testing initiated.  The duty cycle is run for a minimum of 3 minutes.

  - During the duty cycle testing, the technician observes the glass electrode in the fixture for catastrophic failure (example shown in Figure 1 above).

  - If failure is observed, the test is concluded and the glass assembly is labeled Failed.  If no failure is observed, the test is continued in step 4.5.

  - The boost voltage is set to 120% (approximately 152 volts) using the ClearView software and a duty cycle testing initiated.  The duty cycle is run for a minimum of 5 minutes.

  - During the duty cycle testing, the technician observes the glass electrode in the fixture for catastrophic failure (example shown in Figure 1 above).

  - If failure is observed, the test is concluded and the glass assembly is labeled Failed.  If no failure is observed, the test is concluded and the glass assembly is labeled Passed.

The NCR provided MRB instructions for PADT test glass until a total of 10 passing glass assemblies are identified.  As per the MRB instructions, eight (8) pieces were released to MJS Designs, Inc., EPIC’s contract manufacturer, for use in ClearView Device assemblies while two (2) pieces were held for further investigation by the supplier.

NCR 07703 initiated a Corrective and Preventive Action Plan #07704.  This report is intended to support the root cause investigation of the failure.

- Subsequent Failure and Testing

Four of the passed glass assemblies were assembled by MJS Designs, Inc. into electrode assemblies which were subsequently installed into four (4) finished ClearView Devices.  These devices were delivered to EPIC for further evaluation in March 2012.  During initial device testing (ENG-017) catastrophic failures were observed in the glass despite the testing performed according to NCR #07703.  These failures reopened the glass failure investigation initiated by NCR #07703.  An additional test method was developed to further sort the glass assemblies in an effort to identify failure before manufacturing in the electrode assembly.

The final test method is performed on assemblies which passed the original test described in section 4.0 above.  An additional test was performed on any glass assemblies which passed the first test.  This test was performed as follows:

  - Connect the electrode test fixture to the 2.5K Ohm resistor and voltage divider.

  - Connect the oscilloscope to the voltage divider and to ground.  Reference Figure 2 below:

Figure 2

  - Energize the glass electrode.  Record the kV measured for each glass assembly.

Per NCMR 1010A and NCMR 07703 Revision B, two test methods were executed on the remaining glass assemblies from lots 110323-01, 110329-01, and 110628-01.  Only assemblies which passed both test methods were marked Pass.  The test results are located in NCMR 07703 Revision B.  A total of 57 glass assemblies were tested.  34 of the 57 pieces passed the test process and were labeled Passed.  These glass assemblies were released to MJS Designs, Inc. for use in the manufacture of electrode assemblies.  It was later discovered that glass assemblies from lot 110323-01 failed during use in devices and engineering has recommended holding glass from this lot for further investigation.

- Discussion and Next Steps

Four glass assemblies were used in the rework of ClearView Devices SN0986, SN0987, SN0988, and SN0989 by MJS Designs, Inc.  These devices have been through extensive finished device testing without catastrophic failure.  However, the tests have only provided a process for identifying glass assemblies which will fail when energy is applied to the assembly.  The root cause of the failure has not yet been determined.  Therefore, the investigation and CPAP 07704 will remain open. Until the root cause has been established and additional or more appropriate preventive action can be identified, Engineering proposes to implement the test methods identified in this report as required testing of all glass assemblies prior to manufacture of electrode assemblies.  The root cause investigation is currently in progress.  A three month extension will be requested so that the investigation can be completed.  During this extension, work will be done with the supplier to identify the cause and propose appropriate preventive action.  The timeline and schedule of this inspection will be tracked through EPIC’s project management process and reviewed at monthly progress meetings which are documented in meeting minutes.

- Conclusions

The glass assembly test methods will be developed into a standard operating procedure and approved through EPIC’s quality system.  This test method will be performed on 100% of all glass assemblies before manufacturer of the electrode assemblies.  This testing will be implemented prior to the receipt of the next lot of glass assemblies.

- Attachments

Attachment A- NCR #07703 Rev. B

Attachment B- 06717 Rev. X1 Electrode Fixture