Approval Signatures:

| Department | Name | Signature | Date |
| --- | --- | --- | --- |
| Author - ENG | Dan Miller |  |  |

| OPS | Mike Stowell |  |  |
| --- | --- | --- | --- |
| Exec | Nancy Rizzo |  |  |
| Clinical | Patsy Folio |  |  |
| QA | Scott Pletzer |  |  |

- Purpose

This document is intended to provide a test plan for evaluating an experimental ClearView ™ device modified with an additional USB connection.  Current devices use one USB connection to provide communication to both the scanner and camera sections of the device.  The experimental device has a separate USB cable for the scanner and the camera.  This test plan is intended to provide a method for conducting the testing as well as to be used as a formal record of the testing activities.

- Scope

This test plan is intended to be limited to ClearView device SN0981, which has been modified to include a separate USB cable, internally and externally, to provide communication and power between the camera and the computer.  Testing will be performed at the clinical site and on the equipment currently in use at the clinical site.

This testing is intended to be exploratory and to provide data on the performance of a dual USB modification.  This is not a validation test.

- Final Report

# A summary of the test results will be generated and routed for approval.  The summary will follow EPIC’s Engineering Studies procedure, EG-020.  Any required deviations from this written test plan will be documented at the time of deviation.  All deviations will be summarized in the final report.

- Procedure – On Site Testing

  - Shut down the site computer and turn off site device.

  - Remove site device and connect to test laptop.  Startup laptop to provide power to camera in order to maintain operating temperature.  Verify the camera properties tab indicates camera is operating.

  - Install test device SN0981 on cart platform.  Attach an additional USB cable between the computer and the lower USB port on the back panel of the device.  Connect the existing USB cable to the upper USB port on the device.

  - Startup computer and wait for the login screen to appear.  Turn on test device.

Buff, or if necessary clean, glass and probe prior to each run.  Settings and number of calibration runs may be changed as needed.  Record any device or software errors on the Test Record that occur during the following steps:

  - Perform calibration testing per Section A of the Test Record.  This is intended to collect performance data during the warm up period.

  - After 30 minutes since camera power up, perform calibration testing per Section B of the Test Record.

  - Perform duty cycle testing per Section C of the Test Record.

  - Perform calibration testing per Section D of the Test Record.  Select one or more settings and auto-calibrate.

  - Perform probe capture testing per section E of the Test Record.  Capture twenty images of the energized probe.  Cancel before processing.

  - Perform human finger image collection per section F of the Test Record.  Using the duty cycle function in the Device Testing tab, take ten images of fingers without the capacitive barrier and ten images with the capacitive barrier.

  - Shut down the site computer and the test laptop, turn off both devices.

  - Install site device on cart platform.  Startup site computer and wait for the login screen to appear.  Turn on test device.

  - Perform Calibration per clinical study instructions.

- Reference Documents

EG-020, Engineering Studies

- Attachments

Attachment A, Test Record

Attachment A – Test Record

| Section A – Calibration Testing |  |  |  |  |
| --- | --- | --- | --- | --- |
| Time Device Turned On: |  |  |  |  |
| Test Date/Test Start Time: |  | Workstation ID: |  |  |
| Voltage: 110% Exposure Delay: 320 |  | ClearView Version: |  |  |
| B100G61 Time ________ P / F | B100G60 Time ________ P / F |  | B100G59 Time ________ P / F |  |
| B95G61 Time ________ P / F | B95G60 Time ________ P / F |  | B95G59 Time ________ P / F |  |
| B90G61 Time ________ P / F | B90G60 Time ________ P / F |  | B90G59 Time ________ P / F |  |
| Errors: |  |  |  |  |
| Signature/Date |  |  |  |  |

| Section B – Calibration Testing |  |  |  |
| --- | --- | --- | --- |
| B100G61 Time ________ P / F | B100G60 Time ________ P / F | B100G59 Time ________ P / F |  |
| B95G61 Time ________ P / F | B95G60 Time ________ P / F | B95G59 Time ________ P / F |  |
| B90G61 Time ________ P / F | B90G60 Time ________ P / F | B90G59 Time ________ P / F |  |
| Errors: |  |  |  |
| Signature/Date |  |  |  |

Section C – Duty Cycle Testing

Duty Cycle Test – Minimum 100 images Time ________

Errors:

Signature/Date

| Section D – Calibration Testing |  |  |  |
| --- | --- | --- | --- |
| Time ________ P / F | Time ________ P / F | Time ________ P / F |  |
| Errors: |  |  |  |
| Signature/Date |  |  |  |

Section E – Probe Capture Testing

Capture 20 images of energized probe Time ________

Errors:

Signature/Date

Section F – Human Image Collection

Capture 20 human finger images Time ________

Errors:

Signature/Date