Approval Signatures:

Author

| Department | Name | Signature | Date |
| --- | --- | --- | --- |
| IS | Andrew Mason |  |  |

Approvers

| Department | Name | Signature | Date |
| --- | --- | --- | --- |
| Exec | Dr. Nancy Rizzo |  |  |
| QA | Scott Pletzer |  |  |
| OPS | Michael Stowell |  |  |

Revision History

| Revision | Revision Date | Reason for Revision/Change Request | Revised By |
| --- | --- | --- | --- |
| 000 | 24Oct12 | Document creation for the release of version 2.1 | Andy Mason |

- Purpose

The protocol will ask the user to generate scores that will be evaluated against a previously validated external scoring mechanism (i.e., SAS - Statistical Analysis Software).

- Scope

This protocol is intended to be used to validate the appropriate scores are being generated by the algorithms.    This includes all the mathematically generated values produced by the EPIC ClearView algorithm.

- Background

As the company continues to strive to improve the accuracy and effectiveness of the ClearView product it has been noted that revision changes to EPIC ClearView software are not consistently evaluated to ensure that non-algorithm changes have no impact on the performance and output of the software.  The implementation of this new protocol will supplement the current group of protocols used to verify and validate new revisions of ClearView software.

The development of the current algorithm is documented in two engineering studies, ENG-025 which documents the development of the Naïve Bayes and Logistic Regression portions of the algorithm and ENG-026 which documents the development of the EPIC Score portion of the algorithm.  These two development documents were translated into software requirements which were documented in SR-203-01, Algorithm Requirements.  These requirements were then used to develop the ClearView Software as documented in the Analysis Functions- Specifications (SS-203).

- Test Plan / Methodology

The testing will consist of processing three (3) pre-selected unique human sets of images through the ClearView software and then comparing the resultant scores to the result of the same image data being processed through the external SAS (Statistical Analysis Software) application.

- Sample Size / Test Frequency

There will be a total of three (3) complete sets of unique human data which consists of a total sixty (60) energized images (20 for each set).  One (1) calibration set, consisting of ten (10) probe images will be used for each of the three human sets.  The resultant data will be passed through both the ClearView calculation engine and an external calculation engine.  The produced values will be compared against one another to verify they are accurate.  This comparison will be performed once.

One (1) of the original three (3) complete sets of unique human data will be resubmitted through the ClearView software.  The results will be compared against the original run to verify that calculation results are consistently produced.

- Materials / Equipment

One (1) ClearView system consisting of a workstation loaded with ClearView software and a device will be used during the execution of this protocol.

- Data Collection

Data Staging

NOTE: These steps are only complete one time to configure the necessary data.  If the scoring algorithm changes with future ClearView revisions, these steps will be repeated.

- Randomly select three (3) sets of patient images and place them in the following location:

O:\QualitySystems\Quality_System_Documents\Non-quality_system_records\ClearView_Validation\Validation Inputs\Calculation Verification\Patient Images

NOTE:  Place one (1) set of these images in the appropriate subfolder at this location.

Verified		 Not Verified	Non-Conformance Ref. No.

- Randomly select one (1) set of calibration images and place them in the following location:

O:\QualitySystems\Quality_System_Documents\Non-quality_system_records\ClearView_Validation\Validation Inputs\Calculation Verification\Calibration Images

Verified		 Not Verified	Non-Conformance Ref. No.

- Execute steps 1-19 in the Data Collection section below to produce the raw data necessary for the SAS external calculation engine to perform its calculations.

Verified		 Not Verified	Non-Conformance Ref. No.

- Send the raw data to the statistician and have them run the data through the SAS external calculation engine.

Verified		 Not Verified	Non-Conformance Ref. No.

- Processed data (likely in spreadsheet form) is returned to EPIC and stored in the following location

O:\QualitySystems\Quality_System_Documents\Non-quality_system_records\ClearView_Validation\Validation Inputs\Calculation Verification\SAS Results

Verified		 Not Verified	Non-Conformance Ref. No.

Data Collection

- Log into the ClearView system using the “Administrator” login.  It is necessary to use this log-in to allow the manually loading of images.

Verified		 Not Verified	Non-Conformance Ref. No.

- Once logged in, click on the “Calibration” button.

Verified		 Not Verified	Non-Conformance Ref. No.

- In the dialog that appears, right click on each of the image placeholders and select the appropriate image from the following location:

O:\QualitySystems\Quality_System_Documents\Non-quality_system_records\ClearView_Validation\Validation Inputs\Calculation Verification

Verified		 Not Verified	Non-Conformance Ref. No.

- When all ten (10) images are loaded select the “Submit Calibration” button.

Verified		 Not Verified	Non-Conformance Ref. No.

- When the calibration successfully completes, finalize the process without saving the images.

Verified		 Not Verified	Non-Conformance Ref. No.

- Click on the “Manage Patients” button and create a new patient

Verified		 Not Verified	Non-Conformance Ref. No.

- Click on the “Find” button.  When the results appear, attempt to locate a subject with the following attributes:

- First Name = Subject 1

- Last Name = ScoreValidation

- Birthdate = 10/15/1980

- Gender = Female

NOTE:  If a patient with these attributes does not exist, create them using the “New Patient” button at the bottom of the dialog and save the patient.

Verified		 Not Verified	Non-Conformance Ref. No.

- When the patient is either found or created, click the “New Capture” button on the “Patient Demographic” button.

Verified		 Not Verified	Non-Conformance Ref. No.

- Answer “No” to the three questions that appear and click the “OK” button.

Verified		 Not Verified	Non-Conformance Ref. No.

- Right click on each of the image placeholders and select the appropriate image from the following location:

O:\QualitySystems\Quality_System_Documents\Non-quality_system_records\ClearView_Validation\Validation Inputs\Calculation Verification\Patient Images

NOTE:  Once within this folder, look for the subfolder matching your patient’s name to locate images to use.

Verified		 Not Verified	Non-Conformance Ref. No.

- Verify that the angles and center point of all of the loaded images match the values in Attachment C for the subject currently being analyzed.

Verified		 Not Verified	Non-Conformance Ref. No.

- Once all images have been loaded, click the “Submit For Analysis” button.

Verified		 Not Verified	Non-Conformance Ref. No.

- When the submit for analysis has completed, close all resulting windows.

Verified		 Not Verified	Non-Conformance Ref. No.

- Repeat steps 6-13 for each of the following patients:

- First Name = Subject 2

- Last Name = ScoreValidation

- Birthdate = 03/03/1935

- Gender = Male

NOTE:  If a patient with these attributes does not exist, create them using the “New Patient” button at the bottom of the dialog and save the patient.

Verified		 Not Verified	Non-Conformance Ref. No.

- First Name = Subject 3

- Last Name = ScoreValidation

- Birthdate = 07/05/1990

- Gender = Female

NOTE:  If a patient with these attributes does not exist, create them using the “New Patient” button at the bottom of the dialog and save the patient.

Verified		 Not Verified	Non-Conformance Ref. No.

- At the ClearView main screen, click the “Search Capture” button.

Verified		 Not Verified	Non-Conformance Ref. No.

- Type in a last name of “ScoreValidation”.

Verified		 Not Verified	Non-Conformance Ref. No.

- Highlight the most recent run for Subject 1, then click on the “Export Data” button.

Verified		 Not Verified	Non-Conformance Ref. No.

- Using the browse dialog that appears, navigate to the following location:

O:\QualitySystems\Quality_System_Documents\Non-quality_system_records\ClearView_Validation\Validation Inputs\Calculation Verification

NOTE:  Navigate to and select the subfolder for the current version.  If it doesn’t exist, create it.

Verified		 Not Verified	Non-Conformance Ref. No.

- Click the “OK” button to export the data.

Verified		 Not Verified	Non-Conformance Ref. No.

- Repeat steps 16-18 for the remaining subjects.

Verified		 Not Verified	Non-Conformance Ref. No.

- In the “Search Capture” dialog, select Subject 1 and click the “View Capture” button.

Verified		 Not Verified	Non-Conformance Ref. No.

- In the screen that appears click the “Resubmit Images For Analysis” button.

Verified		 Not Verified	Non-Conformance Ref. No.

- When the analysis is complete, close all windows to return to the ClearView main screen.

Verified		 Not Verified	Non-Conformance Ref. No.

- Click on the “Search Capture” button and locate the most recent treatment for Subject 1 (this will be the resubmitted data).

Verified		 Not Verified	Non-Conformance Ref. No.

- Highlight the most recent run for Subject 1, then click on the “Export Data” button.

Verified		 Not Verified	Non-Conformance Ref. No.

- Using the browse dialog that appears, navigate to the following location:

O:\QualitySystems\Quality_System_Documents\Non-quality_system_records\ClearView_Validation\Validation Inputs\Calculation Verification\{Current Version Being Tested}\Resubmitted Data

NOTE:  Navigate to and select the subfolder for the current version.  If it does not exist, create it.

Verified		 Not Verified	Non-Conformance Ref. No.

Data Comparison

- All resultant report values from the SAS external calculation engine will be compared with the equivalent resultant values from the ClearView calculation engine.

Verified		 Not Verified	Non-Conformance Ref. No.

- Acceptance Criteria

All resultant report values from the SAS external calculation engine must be mathematically identically to the equivalent resultant values from the ClearView calculation engine.

- Deviations / Amendments

Any deviations, amendments, or nonconformances that occur during the course of this testing should be recorded on the attached deviations or nonconformances forms (Attachments A & B respectively), and submitted with this protocol at the end of the testing process.

- Attachments

Attachment A – Deviations From Protocol

Attachment B – Nonconformances Worksheet

Attachment C - Center Point & Angle Values

# Attachment A

Deviations from Protocol

NOTE:  This form is used for minor deviations from the protocol as written.  Fill out all sections of this form prior to moving forward in the validation process.

| No. | Date | Name | Description of Deviation (include reference to the protocol section) | Resolution/Action Taken | Reviewed |
| --- | --- | --- | --- | --- | --- |

Attachment B - Non-conformances Worksheet

NOTE:  This form is used for all protocol steps which did not perform as expected.  Fill out all sections of this form prior to moving forward in the validation process.

| No. | Date | Name | Description of Non-conformance (include reference to the protocol section) | Resolution/Action Taken |
| --- | --- | --- | --- | --- |

Attachment C - Center Point & Angle Values

|  | Subject 1 |  |  | Subject 2 |  |  | Subject 3 |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Finger Image | X | Y | Angle | X | Y | Angle | X | Y | Angle |
| Right Thumb - Unfiltered | 162 | 111 | -1 | 158 | 113 | 1 | 154 | 110 | 6 |
| Right Index - Unfiltered | 168 | 107 | 4 | 157 | 106 | 7 | 139 | 105 | 4 |
| Right Middle - Unfiltered | 168 | 105 | -1 | 163 | 105 | 7 | 158 | 103 | 2 |
| Right Ring - Unfiltered | 144 | 108 | 2 | 151 | 104 | 6 | 163 | 98 | 2 |
| Right Pinky - Unfiltered | 167 | 103 | 4 | 152 | 109 | 11 | 153 | 108 | 18 |
| Left Thumb - Unfiltered | 141 | 122 | -1 | 150 | 116 | 2 | 153 | 117 | 6 |
| Left Index - Unfiltered | 152 | 112 | -2 | 157 | 109 | 3 | 156 | 109 | 1 |
| Left Middle - Unfiltered | 150 | 109 | 0 | 161 | 108 | -6 | 153 | 103 | -7 |
| Left Ring - Unfiltered | 155 | 105 | -11 | 154 | 107 | -8 | 134 | 107 | -11 |
| Left Pinky - Unfiltered | 170 | 103 | -10 | 137 | 110 | -18 | 150 | 111 | -13 |
| Right Thumb - Filtered | 170 | 113 | 12 | 157 | 110 | 1 | 148 | 106 | 2 |
| Right Index - Filtered | 162 | 106 | 8 | 166 | 105 | 6 | 155 | 106 | 10 |
| Right Middle - Filtered | 160 | 99 | 7 | 139 | 121 | -4 | 164 | 103 | 10 |
| Right Ring - Filtered | 162 | 102 | 11 | 144 | 101 | -2 | 157 | 110 | 3 |
| Right Pinky - Filtered | 168 | 106 | 14 | 147 | 107 | 12 | 134 | 111 | 14 |
| Left Thumb - Filtered | 142 | 121 | 1 | 142 | 124 | -7 | 150 | 121 | 7 |
| Left Index - Filtered | 159 | 104 | -4 | 148 | 106 | 0 | 168 | 111 | 3 |
| Left Middle - Filtered | 141 | 109 | -5 | 158 | 111 | -8 | 149 | 105 | -1 |
| Left Ring - Filtered | 137 | 122 | -15 | 150 | 98 | -8 | 143 | 104 | -11 |
| Left Pinky - Filtered | 140 | 109 | -11 | 145 | 110 | -17 | 151 | 106 | -12 |