ORIGINATOR:

| Name | Job Title | Signature | Date |
| --- | --- | --- | --- |
| Brian Cullinan | Software Developer |  |  |

APPROVALS:

| Name | Job Title | Signature | Date |
| --- | --- | --- | --- |
| Nancy Rizzo | Founder, President and CTO |  |  |
| Lloyd Kurth | Technology Director |  |  |
| Dan Miller | Manufacturing Engineer |  |  |
| Jan Ayres | Operations Director |  |  |
| Andrea Miller | Quality Assurance Director |  |  |

# Purpose

The purpose of this document is to define the software requirements for the ClearView software.

# Scope

This document describes the ClearView software functional requirements. Algorithm software requirements are described in a separate document, SR-207, ClearView Software Requirements, Algorithm.

This document does not describe the hardware requirements and overall system requirements for ClearView, which are documented in DD-032, Product Requirements ClearView System. The ClearView software requirements are subordinate to the system requirements. Traceability to specific system requirements is provided in Appendix A.

# Definitions

| Authentication | The process of identifying the user accessing a computer or computer application. |
| --- | --- |
| Authorization | The process of granting access to application functions based on the identity of the authenticated user. |
| Capture | A capture is a single image taken from a patient. |
| Database | A comprehensive collection of related data organized for convenient access. Examples include, but are not limited to, relational databases, object databases, lightweight loosely structured databases (such as NoSQL), and a simple file system. |
| Framework | A set of industry standard tools that is available with the platform to help solve the task at hand. |
| Image | A single frame captured from the camera of the imaging device. |
| Navigation | The ability for the user to find and operate the functions of a software application. |
| Privileged access | Privileged access will be used to denote functions that are restricted to use by EPIC staff. |
| Scan | A scan is a set of 20 finger images captured from a patient. |
| Setup Verification | The process of determining whether or not the imaging device produces images within tolerance. |
| User | A user is the individual using the application. This does NOT refer to any specific user or role that is defined. The use of the word “user” in this document does not indicate access restrictions and/or software permissions. |

# Requirements

The requirements described below define how the user controls the functions of the ClearView device through the software interface/GUI.

General

ClearView software shall clearly indicate its current version, which may include other version information (e.g., library version information, etc.) to describe software components.

ClearView software shall store an audit trail for all changes made to the database. The audit trail shall contain:

Changed data

Date and time of the change

Currently logged in user when the change occurred

ClearView may contain a main menu or home screen for easy navigation.

Navigation controls shall be provided to access the functions and guide the user through the order of operations listed in the following sections.

All necessary data shall be persisted to a local database for easy retrieval by the application and also in human readable form.

Proper security mechanisms shall be used to secure the application from unauthorized use and to comply with HIPAA regulations. Such mechanisms include, but are not limited to:

Require a valid username and password.

Close the application after a specified amount of time of inactivity.

Prevent the user from entering passwords too quickly in an attempt to brute-force the system, also known as “throttling”.

Check for proper authentication and authorization throughout important system functions such as changing camera settings or functions requiring privileged access.

Code obfuscation to prevent the decompiling and interpreting of internal methods and objects.

The software shall establish a single connection to communicate with the firmware in the device in order to control device functions, reconnecting as needed, to maintain the connection.

The software shall establish a single connection to the camera in order to control the camera, reconnecting as needed, to maintain the connection.

Device status, limited to connected or disconnected, shall be displayed to the user.

Users with privileged access should be able to:

View, clear and save detailed device information provided by the firmware.

Capture and save a single image.

Licensing

A “license” may be used to control what functionality is available for a set of roles and/or users.

The active license should be clearly indicated.

Users should not be able to modify the license.

ClearView software should produce detailed levels of logging that can be used to debug the system. Device events and commands should be included in these logs.

The ClearView software shall meet the appropriate regulatory requirements detailed by the following documents:

21 CFR Part 11 FDA Electronic Records; Electronic Signatures

21 CFR 820 FDA Quality System Regulation

Health Insurance Portability and Accountability Act of 1996 (HIPAA)

ISO 13485 Medical devices – Quality management systems – Requirements for regulatory purposes

Users shall be able to cancel out of all functions unless otherwise detailed in the requirements.

Users

A user shall be required to log in with a unique, valid username and password before using the ClearView software.

New users can be added. When a user is added, at least the following information shall be required:

Name (First and Last Name)

User email address

Unique username

A password which meets the following requirements:

Passwords shall implement aging to prevent the same password being used.

Passwords shall implement a specified minimum length to ensure it is secure.

Passwords shall use a salted one way hash to ensure reasonable difficulty against a brute force attack.

The user shall be notified of any invalid entry.

An existing user shall be able to modify his/her own:

Name

Password (such that it still fulfills requirements SR-2.2.4)

The user shall be notified of any invalid entry.

A user shall be assigned to one or more roles in order to control what functionality they have access to.

The following roles should be defined:

Administrator – Ultimate power over all functionality of the system and privileged functionality which is intended to only be used by EPIC staff.

Tech administrator – The ability to load and export finger images, and set brightness and gain controls.

Tech user – Can perform basic setup verification and capture functions.

Additional roles may be defined and must be documented in the product specifications.

Each role may grant access to specific functionality, such as access to press a button which is otherwise hidden if the user does not have access. This level of granularity is called “permissions” and it may be assigned to a specific role.

The application shall allow the logged in user to be switched easily and, when possible, place the new user in the same state as the previous user.

An existing user can be deactivated. A deactivated user shall not be able to log in to use the ClearView software.

Patients

It shall be possible to add a new patient.

Checks shall be implemented to reasonably prevent duplicate patients from being added to the database. The software shall notify users if a potential duplication is identified.

Existing patients should be searchable, and search results sortable, on at least the following fields:

Name (i.e., first name, last name, and middle initial)

Gender

Birthdate

The following minimum information should be stored about each patient:

Name (i.e., first name, last name, and middle initial)

Birthdate

Gender

It shall be possible to view and modify all patient information.

The software shall notify the user and require the corresponding corrections to be made to invalid or missing information.

Setup Verification

On user demand, the setup verification process shall be able to automatically capture a set of images to be used with the algorithm.

A setup verification which passes the analysis criteria should be performed before a scan if any of the following conditions are met:

A setup verification has not occurred within reasonable intervals to ensure accurate setup.

Any of the camera settings (i.e., exposure delay, brightness, gain, exposure duration and percent of power) have been changed.

The ClearView device USB connection for communicating with the firmware is reconnected after being disconnected.

The ClearView computer has been restarted (rebooted).

The setup verification process shall indicate progress and if the setup verification has passed or failed.

The setup verification process passes if all images in the set pass. The criteria used to determine a passing or failing image is detailed in SR-207, ClearView Software Requirements, Algorithm.

The setup verification process shall capture and not store 5 warm-up images.

The camera connecting to the software should be configurable.

Privileged access should allow the configuration of multiple sets of predefined device controls for at least the following:

Brightness

Gain

Exposure duration

Exposure delay

Percent of power (boost voltage)

Any user shall be able to choose one of the predefined sets of device controls defined in SR-4.7 before the automatic setup verification process begins. A default pre-selected set may be identified.

The live preview and the captured images shall be displayed.

Successful setup verification (including images and calculated numeric data) shall be stored in the database with the current date and time.

Previous successful setup verifications should be viewable on demand.

Setup verification data should be exportable, including images and all calculated numeric data.

Failed setup verification including the images collected and the calculated numeric data should be viewable and exportable, by a user with privileged access, at the time of failure.

The setup verification process should detect failed images given a set of adjustable parameters such as:

Number of bright pixels per region

Average intensity standard deviations per region

Number of acceptable failures

Users with privileged access should be able to load images into the setup verification analysis routine manually.

Image Capture

Scans should be searchable and sortable on at least the following fields:

Patient name

Scan date

Software version

Indication of resubmitted data

The list of scans shall default to sorting by newest to oldest.

Stored images (including the setup verification associated with the scan) shall be viewable. That is, the user can view all related images in a similar display used to capture patient images.

A patient name shall be selected before starting a scan for use with image analysis. The selected patient name will be displayed during image capture.

A scan shall be associated with the most recent passing setup verification record to be used with image analysis.

A scan should allow for the recording of extra information and demographics about the patient before a scan is initiated. This information includes, but is not limited to the following questions/data requirements:

Has the patient consumed wheat in the last 2 hours? Yes or No

Has the patient consumed alcohol in the last 24 hours? Yes or No

Has the patient consumed caffeine in the last 2 hours? Yes or No

User shall be notified if these questions are not completed.

Scans shall be stored in the database.

Image Collection

A user must be able to collect a new set of images for a patient. Images shall be associated with the patient data in the database.

All images and alignment values should be stored in the database, including the finger image used to align the sectors properly, with the date and time of the capture.

A live preview of the finger before the energized image is captured shall be displayed.

Instructions and status of current finger may be displayed to the user throughout the capture process to ensure the systematic capture of all 10 finger images with and without a capacitive barrier.

The energized image and the captured finger image used for image alignment shall be displayed immediately after capture.

ClearView software should automatically set the current angle, X and Y positions at the start of collection of an image.

ClearView software shall allow the user to modify/correct the image alignment.

Image alignment should be displayed on both the energized image and the finger image.

Image alignment may be displayed by:

An ellipse which is defined by the inner ring of the energized image.

A crosshair (two mutually perpendicular fine lines) that places the X/Y intercept at the center of ellipse.

The Y axis aligns to the centerline of the finger and defines the current angle. This axis should extend below the ellipse.

The X axis should extend slightly beyond the ellipse on each side.

The actual angle and midpoint values should display numerically.

The user should be able to turn the image alignment display on and off.

The user should be able to recapture any image at any point during image capture.

Users with privileged access should be able to load images for analysis manually.

Users with privileged access should be able to export both the energized and finger images to the file system before analysis begins.

Users shall be prompted when collection is complete to ensure that all images and image alignment is correct before analysis begins.

Users shall be able to initiate analysis after all 20 images have been collected.

Submit for Analysis

When captured images are submitted for analysis, the software version used to generate the data shall be recorded.

The date of original image capture shall be recorded.

All image collection requirements defined in section SR-5.8 must be met at the time of submission.

Resubmit for Analysis

Users with privileged access shall be able to resubmit images for analysis.

When images are resubmitted for analysis, all subsequent reports and data displayed shall clearly denote that it has been resubmitted from previously collected data.

All requirements detailed in sections SR-5.8 and SR-5.9 should be met at the time of analysis.

All data should be exportable by users with privileged access. Capture information includes, but is not limited to:

Patient data (de-identified data only as defined in US Title 45 section 164.514a)

Setup verification images and data

Images, including alignment values (center-point and angle)

Coefficient data

Reported data

Any exported information should be saved to the location on disk selected by the user.

Reports

Reports should be customizable/interchangeable based on a “license” or specific permission.

Reports shall be displayed to the user based on license and permission.

Users with privileged access should be able to view the following additional data:

Raw data

Data used in generating a report

All full-sized images (i.e., setup verification, energized and finger)

Sector divisions

Reports should include:

Date of the scan

Software version used to generate the report

Patient biographical data

Reported data from all organ/structures defined in Appendix A, with physical and autonomic values

Reports shall be printable, and secured (non-modifiable) digital copies may be available on demand.

Reports shall be searchable on at least the following fields:

Date of the scan

Software version used to generate the report

Patient biographical data

Reports may be accessible during patient and scan searches.

Users should be able to view a graphical representation of the NS coefficient described in SR-207, ClearView Software Requirements, Algorithm.

Users should be able to view a graphical representation of energized finger images mapped to body locations; this is known as “biofield”. A magnified view of the graphical representation should be displayed.

Users with privileged access should be able to see sector alignment with the biofield display.

Energized images should be flattened and displayed to the user.

# References

|  | 21 CFR Part 11, FDA Electronic Records; Electronic Signatures |
| --- | --- |
|  | 21 CFR 820, FDA Quality System Regulation |
|  | Health Insurance Portability and Accountability Act of 1996 |
|  | ISO 13485, Medical devices – Quality management systems – Requirements for regulatory purposes |
| DD-032 | Product Requirements ClearView System |
| SR-207 | ClearView Software Requirements, Algorithm |

# Appendices

Appendix A: ClearView Organs/Structures

Appendix B: Traceability to ClearView System Requirements

Appendix C: Traceability to CFR 21 Part 11

ClearView Organs/Structures

Eye (L)

Eye (R)

Ear/Nose/Sinus (L)

Ear/Nose/Sinus (R)

Jaw/Teeth (L)

Jaw/Teeth (R)

Cervical Spine

Thoracic Spine

Lumbar Spine

Sacrum

Coccyx/Pelvis

Nervous System

Hypothalamus

Pituitary

Pineal

Cerebral Cortex

Cerebral Vessels

Immune System

Spleen

Respiratory System

Respiratory/Mammary

Thorax Respiratory

Throat/Larynx/Trachea/Thyroid

Coronary Vessels

Heart

Heart (Left Side)

Heart (Right Side)

Cardiovascular Circulation

Duodenum

Jejunum

Ileum

Blind Gut

Abdominal Region

Thyroid

Pancreas

Liver

Gallbladder

Appendix

Ascending Colon

Transverse Colon

Descending Colon

Sigmoid Colon

Rectum

Renal & Reproductive

Adrenal

Kidney - Left

Kidney - Right

Urokidney

Genitourinary System

Traceability to ClearView System Requirements

The following matrix identifies which software requirements address each product (system) requirement.

| Product Requirement | Product Requirement Text | Software Requirements |
| --- | --- | --- |
| PR-1.5 | System shall be used by trained personnel. | SR-1.6 SR-1.6.1 SR-1.6.4 |
| PR-1.7 | System operation and output should be efficient and clear. | SR-1.3 SR-1.4 SR-1.14 SR-2.8 |
| PR-2.1 | Software shall control device functions. | SR-1.7 SR-1.8 SR-4.1 SR-4.5 SR-4.6 SR-4.7 SR-4.8 SR-5.8.1 SR-5.8.11 |
| PR-2.2 | Software shall prompt for, collect and save setup verification images. | SR-1.5 SR-4.1 SR-4.2 (all) SR-4.5 SR-4.9 SR-4.10 SR-4.12 SR-4.13 SR-4.15 |
| PR-2.3 | Software shall perform analysis on setup verification images. | SR-4.3 SR-4.4 SR-4.14 (all) |
| PR-2.4 | Software shall prompt for, collect and save patient information. | SR-1.5 SR-3.1 SR-3.2 SR-3.4 (all) SR-3.5 SR-3.6 SR-5.6 (all) |
| PR-2.5 | Software shall prompt for, collect and save patient finger images. | SR-1.5 SR-5.4 SR-5.7 SR-5.8 (all) |
| PR-2.6 | Software shall perform analysis on patient finger images. | SR-5.5 SR-5.9 SR-5.9.1 SR-5.9.3 |
| PR-2.7 | Software shall generate and display report of measurements. | SR-6.2 SR-6.3 (all) SR-6.4 (all) SR-6.7 SR-6.8 SR-6.9 SR-6.10 SR-6.11 |
| PR-2.8 | Software shall generate printable report of measurements. | SR-6.5 |
| PR-2.9 | Software shall have search, view and edit functions for patient information. | SR-1.5 SR-3.3 (all) |
| PR-2.10 | Software shall have search and view functions for images, scans and reports. | SR-1.5 SR-4.11 SR-5.1 (all) SR-5.2 SR-5.3 SR-5.7 SR-5.9.1 SR-5.9.2 SR-5.10.2 SR-6.2 SR-6.3 SR-6.6 (all) SR-6.7 |
| PR-2.11 | Software shall have separate privileged and non-privileged user functions. | SR-1.6.4 SR-1.10 (all) SR-2.4 SR-2.5 (all) SR-2.6 SR-2.7 |
| PR-2.12 | Software shall allow users with privileged access to modify image alignment and resubmit patient scan for analysis. | SR-1.6.4 SR-5.10 SR-5.10.1 SR-5.10.3 |
| PR-2.13 | Software shall allow users with privileged access to export patient scan information. | SR-1.6.4 SR-5.8.13 SR-5.11 (all) SR-5.12 |
| PR-2.14 | Software shall prohibit users from access to code. | SR-1.6.5 |
| PR-2.15 | Software shall communicate with the device firmware to limit device operation to appropriate levels of power, frequency and duration. | SR-1.7 SR-1.10 (all) SR-4.1 SR-4.5 SR-4.7 (all) SR-4.8 SR-5.8.1 SR-5.8.11 |
| PR-2.16 | Software may use licensing to allow report customization. | SR-1.11 (all) SR-6.1 SR-6.2 SR-6.4.4 |
| PR-2.17 | Software may include functionality to aid in investigation and correction of system issues. | SR-1.1 SR-1.12 |
| PR-3.2 | The camera shall be controllable by software. | SR-1.8 SR-1.9 SR-4.1 SR-4.5 SR-4.6 SR-4.7 (all) SR-4.8 SR-5.8.1 SR-5.8.11 |
| PR-6.1 | System shall comply with the regulations for saleable devices in the intended markets, including but not limited to the United States and Europe. | SR-1.2 (all) SR-1.6 SR-1.6.1 SR-1.6.2 SR-1.6.3 SR-1.13 (all) SR-2.1 SR-2.2 (all) SR-2.3 (all) SR-2.4 SR-2.9 |

Traceability to CFR 21 Part 11

This table defines which ClearView requirements enforce Part 11 compliance as well as any related quality system procedures that ensure the Part 11 requirement is met.

| Part 11 Requirement | ClearView Requirement |
| --- | --- |
| Subpart B Sec. 11.10 (Closed System) |  |
| (a) Validation of systems to ensure accuracy, reliability, consistent intended performance, and the ability to discern invalid or altered records. | EG-001, Design Control Procedure |
| (b) The ability to generate accurate and complete copies of records in both human readable and electronic form suitable for inspection, review, and copying by the agency. Persons should contact the agency if there are any questions regarding the ability of the agency to perform such review and copying of the electronic records. | SR-1.5, SR-6.2, SR-6.3, SR-6.4, SR-6.5 |
| (c) Protection of records to enable their accurate and ready retrieval throughout the records retention period. | SR-1.2, SR-1.6, SR-2.1, SR-6.5 |
| (d) Limiting system access to authorized individuals. | SR-1.6, SR-2.1 |
| (e) Use of secure, computer-generated, time-stamped audit trails to independently record the date and time of operator entries and actions that create, modify, or delete electronic records. Record changes shall not obscure previously recorded information. Such audit trail documentation shall be retained for a period at least as long as that required for the subject electronic records and shall be available for agency review and copying. | SR-1.2 |
| (f) Use of operational system checks to enforce permitted sequencing of steps and events, as appropriate. | SR-1.4, SR-1.14, SR-1.6, SR-2.2.5, SR-2.3.3, SR-3.2, SR-3.6, SR-5.8.3, SR-5.8.14 |
| (g) Use of authority checks to ensure that only authorized individuals can use the system, electronically sign a record, access the operation or computer system input or output device, alter a record, or perform the operation at hand. | SR-1.6, SR-1.10 (all), SR-1.11 (all), SR-2.1, SR-4.7 (all), SR-4.13, SR-4.15, SR-5.8.12, SR-5.8.13, SR-5.10.1, SR-5.11 (all), SR-6.3, SR-6.10 |
| (h) Use of device (e.g., terminal) checks to determine, as appropriate, the validity of the source of data input or operational instruction. | Clinical Procedures |
| (i) Determination that persons who develop, maintain, or use electronic record/electronic signature systems have the education, training, and experience to perform their assigned tasks. | Training Procedures |
| (j) The establishment of, and adherence to, written policies that hold individuals accountable and responsible for actions initiated under their electronic signatures, in order to deter record and signature falsification. | EG-001, Design Control Procedure |
| (k) Use of appropriate controls over systems documentation including: (1) Adequate controls over the distribution of, access to, and use of documentation for system operation and maintenance. (2) Revision and change control procedures to maintain an audit trail that documents time-sequenced development and modification of systems documentation. | EG-001, Design Control Procedure |
| Sec. 11.50 Signature manifestations. |  |
| (a) Signed electronic records shall contain information associated with the signing that clearly indicates all of the following: (1) The printed name of the signer; (2) The date and time when the signature was executed; and (3) The meaning (such as review, approval, responsibility, or authorship) associated with the signature. | The ClearView Software is expected to be classified as a Class II non-invasive device. As such, the operator is not required to document a signature when data is collected. The ClearView Software will not be developed to be compliant with any portion of this section. |
| (b) The items identified in paragraphs (a)(1), (a)(2), and (a)(3) of this section shall be subject to the same controls as for electronic records and shall be included as part of any human readable form of the electronic record (such as electronic display or printout). |  |
| Sec. 11.70 Signature/record linking. |  |
| Electronic signatures and handwritten signatures executed to electronic records shall be linked to their respective electronic records to ensure that the signatures cannot be excised, copied, or otherwise transferred to falsify an electronic record by ordinary means. | See section 11.50 above. |
| Sec. 11.100 General requirements. |  |
| (a) Each electronic signature shall be unique to one individual and shall not be reused by, or reassigned to, anyone else. | See section 11.50 above. |
| (b) Before an organization establishes, assigns, certifies, or otherwise sanctions an individual`s electronic signature, or any element of such electronic signature, the organization shall verify the identity of the individual. |  |
| (c) Persons using electronic signatures shall, prior to or at the time of such use, certify to the agency that the electronic signatures in their system, used on or after August 20, 1997, are intended to be the legally binding equivalent of traditional handwritten signatures. |  |
| (1) The certification shall be submitted in paper form and signed with a traditional handwritten signature, to the Office of Regional Operations, 12420 Parklawn Drive, RM 3007 Rockville, MD 20857. |  |
| Sec. 11.200 Electronic signature components and controls. |  |
| (a) Electronic signatures that are not based upon biometrics shall: (1) Employ at least two distinct identification components such as an identification code and password. | See section 11.50 above. |
| (i) When an individual executes a series of signings during a single, continuous period of controlled system access, the first signing shall be executed using all electronic signature components; subsequent signings shall be executed using at least one electronic signature component that is only executable by, and designed to be used only by, the individual. |  |
| (ii) When an individual executes one or more signings not performed during a single, continuous period of controlled system access, each signing shall be executed using all of the electronic signature components. |  |
| (2) Be used only by their genuine owners; and |  |
| (3) Be administered and executed to ensure that attempted use of an individual's electronic signature by anyone other than its genuine owner requires collaboration of two or more individuals. |  |
| (b) Electronic signatures based upon biometrics shall be designed to ensure that they cannot be used by anyone other than their genuine owners. |  |
| Sec. 11.300 Controls for identification codes/passwords. |  |
| (a) Maintaining the uniqueness of each combined identification code and password, such that no two individuals have the same combination of identification code and password. | SR-2.2.3, SR-2.2.4 (all) EPIC will verify and validate all password and database security methods to ensure that they function properly and work as intended. Customers will have the responsibility to ensure that security measures are used properly to prevent inappropriate access and system use. |
| (b) Ensuring that identification code and password issuances are periodically checked, recalled, or revised (e.g., to cover such events as password aging). | SR-2.2.4.1 EPIC will verify and validate all password and database security methods to ensure that they function properly and work as intended. Customers will have the responsibility to ensure that security measures are used properly to prevent inappropriate access and system use. |
| (c) Following loss management procedures to electronically deauthorize lost, stolen, missing, or otherwise potentially compromised tokens, cards, and other devices that bear or generate identification code or password information, and to issue temporary or permanent replacements using suitable, rigorous controls. | SR-2.3, SR-2.3.2, SR-2.9 EPIC will verify and validate all password and database security methods to ensure that they function properly and work as intended. Customers will have the responsibility to ensure that security measures are used properly to prevent inappropriate access and system use. |
| (d) Use of transaction safeguards to prevent unauthorized use of passwords and/or identification codes, and to detect and report in an immediate and urgent manner any attempts at their unauthorized use to the system security unit, and, as appropriate, to organizational management. | SR-1.6.3, SR-1.6.4, SR-2.1 EPIC will verify and validate all password and database security methods to ensure that they function properly and work as intended. Customers will have the responsibility to ensure that security measures are used properly to prevent inappropriate access and system use. |
| (e) Initial and periodic testing of devices, such as tokens or cards, that bear or generate identification code or password information to ensure that they function properly and have not been altered in an unauthorized manner. | SR-1.6.4, EPIC will verify and validate all password and database security methods to ensure that they function properly and work as intended. Customers will have the responsibility to ensure that they are not altered during system use. |

Revision History

| Revision | Approval Date | Reason for Revision/Change Request | Revised By |
| --- | --- | --- | --- |
| 000 |  | Update the software requirements to combine them into one document referencing the algorithm requirements document. This new document obsoletes the following: SR-001, Patient Demographic Functions Requirements SR-202, Capture Functions Requirements SR-202-01, Scan Resubmit Requirements SR-202-02, Persisting Finger Image Requirements SR-202-03, Quick Export Requirements SR-203, Analysis Function Requirements SR-204, Camera Functions Requirements SR-205, Application Functions Requirements | B. Cullinan |