ORIGINATOR:

| Name | Job Title | Signature | Date |
| --- | --- | --- | --- |
| Brian Cullinan | Software Developer |  |  |

APPROVALS:

| Name | Job Title | Signature | Date |
| --- | --- | --- | --- |
| Nancy Rizzo | Founder, President and CTO |  |  |
| Lloyd Kurth | Technology Director |  |  |
| Dan Miller | Manufacturing Engineer |  |  |
| Jan Ayres | Operations Director |  |  |
| Andrea Miller | Quality Assurance Director |  |  |

# Purpose

The purpose of this document is to describe the requirements for the image analysis process. This document is supplementary to SR-206, ClearView Software Requirements.

# Scope

This document consists of analysis requirements for setup verification images as well as energized finger images.

# Definitions

| Bitmap | A graphic format used by the system. All images captured by the system are stored in this format. The images are all 320 pixels by 240 pixels in size. |
| --- | --- |
| Coefficients | Numerical calculations that describe specific attributes of the image. |
| Encryption | Modifying human readable text into a format that is not readable without performing a conversion. |
| Energized Image | The image captured by the scanner showing the release of energy. This image is captured and stored as a bitmap having a width of 320 pixels and a height of 240 pixels and 24 bit color depth. An example of an energized image is show below: |
| Filtered Image | An energized image obtained by placing a specially designed filter between the subject’s finger and the glass electrode. |
| Finger Image | The raw image of the finger captured before the voltage is applied to the glass electrode. |
| Finger Sector | A portion of the energized image that is analyzed for calculating the various coefficients. Each of the finger images has between 6 and 9 sectors for calculation purposes. The sectors are represented as 'pie slices' emanating from the calculated center of the image. An example of these sectors is displayed below: |
| Fractal | A geometrical or physical structure having an irregular or fragmented shape. In an energized image, the fractal coefficient identifies the presence of a repeated pattern in the image. |
| Naïve Bayes | A naive Bayes classifier is a simple probabilistic classifier based on applying Bayes' theorem with strong (naive) independence assumptions. A more descriptive term for the underlying probability model would be "independent feature model". |
| Noise Level | This is the pixel intensity (a value between 0 and 255) that represents the point at which we will consider the pixel “on”. |
| Pixel | A graphical component that has the ability to turn on and off and display color and intensity with values ranging from 0 to 255. |
| Plasma | Plasma is one of the four fundamental states of matter which occurs by heating gas which ionizes molecules turning it into plasma, which contains charged particles: positive ions and negative electrons or ions. |
| Scan | The process of collecting the energized images from the fingers of a subject. A complete scan will consist of 10 filtered and 10 unfiltered images. |
| Sector Concentric Ring | Each of the finger sectors is further divided in to 4 equal size radii illustrated below. |
| Setup Verification Image | An image captured using the calibration probe rather than a human finger. This is used as a baseline in performing some image calculations. |
| Standard Deviation | A measure of the variance or dispersion from the average (mean), or expected value. |
| Unfiltered Image | An energized image obtained by placing the subject’s finger directly on the glass electrode. |
| Z-Score | In statistics, a standard score indicated by how many standard deviations an observation or datum is above or below the mean. It is a dimensionless quantity derived by subtracting the population mean from an individual raw score and then dividing the difference by the population standard deviation. |

# Requirements

General

Technologies focused on processing mathematical functions should be used.

Design specifications should define all mathematical functions in human readable form.

Energized images shall be bitmaps, 320 pixels wide by 240 pixels high with 24 bit color depth.

Setup Verification Analysis

The setup verification analysis shall evaluate setup verification images.

A setup verification image should be divided into 4 sections for analysis. The sections are defined as:

Inner, the inside circle of the corona (should be mostly black)

Outer, the outside space not including the corona

Corona, the dimmer reaching arms of the plasma

High plateau, the bright edge between the inner ring and where the pixels begin to fade in to distinguishable “fingers”

The following characteristics should be determined for each section (defined in AR-2.2) of the setup verification image.

Area (in pixels)of the corona

Brightness of pixels (shape of histogram)

Amount of noise (pixels should be nearly black outside the corona)

The characteristics of the setup verification image shall be compared to an acceptance criterion.  Any values outside of the acceptance criteria shall fail the analysis.  Outputs from this comparison should include:

Number of pixels that appear outside the acceptance criteria for each image

The characteristics of the setup verification image and the acceptance criteria for each pixel are compared.  The sum of the differences of this comparison is reported for each section of each image.

Scan Coefficients

The algorithm should capture the following attributes of each energized finger image for use with the EPIC Scoring Algorithm:

The current set of passing setup verification images will be analyzed to determine an acceptable noise level. This noise level is used to determine some of the following coefficients.

Energized finger images shall be divided into sectors that are predefined and correlate to one or more organ components.

The sector area will be determined by adding up all the bright pixels above the noise level.

The area should be normalized by dividing sector area by the number of sectors and the angle.

An average intensity value for the sector should be determined by averaging the pixels above the noise level.

An entropy (amount of disorder) value should be measured radially from the center using ¼ of a degree for each angle within the sector.

Specific attributes of an image should be measured:

Amount of change in pixel intensity along the degree line

Gaps around the edge of the energized ellipses

Abrupt breaks (sudden drop of intensity)

An overall form value should be determined by weighting and summing the attributes described in AR-3.7.

Additional ring form values should be determined for each of the four concentric rings for each sector by weighting and summing the attributes described in AR-3.7.

A secondary form value should be determined by weighting and summing the attributes in AR-3.7 for the first two concentric rings.

A third form value should be determined similar to AR-3.10 except using a different method of traversal.

A break coefficient should be determined for each sector measuring breaks (sudden drop in pixel intensity).

A fractal value calculated using standard boxing for a two dimensional area should be determined.

An NS value should be calculated by comparing the energized finger image with the associated setup verification image.

A ring thickness value for the entire corona of the energized finger image should be calculated as the radial distance from the innermost edge of the corona to the outermost edge of the corona, for each sector.

A ring intensity value should be calculated as the average pixel intensity for the radial distance from the innermost edge of the corona, to the outermost edge of the corona, for each sector.

Four average intensity values should be calculated by averaging all pixels within each of the four concentric rings for each sector.

Scoring Algorithm

- The following additional values will be computed as a part of the EPIC Score which summarizes and prioritizes the coefficient values:

The standard z-score should be calculated across all coefficient values in a given sector.

The base score for each coefficient will be calculated for each finger sector. This score is a product of the weighting factor (which is unique to each coefficient and will be detailed in the specifications), the calculated coefficient and the coefficient z-score.

The total score is a sum of the individual base scores for each sector.

Specific rules should be used to rank the coefficients specified in section AR-3. These rules combined with the total score create a final score reported to the user.

Logistic regression should be calculated for each of the five organ systems using the calculated scan coefficients.

Odds ratio should be calculated using coefficient data.

The reported score should be determined by taking into account the EPIC score and the logistic regression.

# References

| SR-206 | ClearView Software Requirements |
| --- | --- |

Revision History

| Revision | Revision Date | Reason for Revision/Change Request | Revised By |
| --- | --- | --- | --- |
| 000 |  | Initial Release Obsoletes SR-203.004, Analysis Function Requirements and portions of SR-204.006, Camera Functions Requirements and SR-202.004, Capture Functions Requirements. | B. Cullinan |