## ClearView Requirements
This table provides the traceability and justification for modifications to the software requirements documents. Historical software requirements were provided in separate documents based on process flow within the software. A decision was made to move to a consolidated version of software requirements into two documents, SR-206 which includes all general software requirements and SR-207 which includes all algorithm specific requirements. This table is a one time document which is being collated to justify the change in documentation. None of the proposed changes actually alter the software requirements from the original approved requirements, unless specifically stated in the table below.
SR-001 rev. 002 Patient Demographics Requirements |  | SR-206 rev. 000 New Requirement (Unless otherwise specified)
Reference | Description | New Reference | Justification
4.1 | Users must be able to search for a patient in the system. All new patient entry must begin with the search screen. This is to ensure that the specific patient being entered does not already exist. | SR-3.2 SR-3.3 | The detail regarding where new patient entry begins will be included in the specifications.
4.2 | To search a patient, the user is prompted to enter at a minimum the patient’s first or last name. Optionally, they can enter the patient’s birth date and gender in order to narrow the search results. | SR-3.3 | Equivalent wording used.
4.3 | The system displays zero or more results for the search criteria entered. | N/A | This detail will be included in the specifications.
4.4 | If the appropriate patient is not found, then the user is presented with the ability to create a new patient profile. | N/A | This detail will be included in the specifications.
4.5 | If the patient is found, then the user is able to view and edit the patient demographics information, capture a new scan for the patient selected, or enter a new patient profile. | SR-3.5 SR-5.4 | Equivalent wording used.
4.6 | Patient demographic fields should be audited. An audit trail should be kept for each data point, consisting of the following tracking points:  ID of the user that performed the change.  Date and time of the change down to the minute.  The before and after values.  The name of the data point. | SR-1.2 | Equivalent wording used.
4.7 | After a search, if the patient is found, then the user is able to view the patient demographics information, capture a new scan for the patient selected, or enter a new patient profile. | SR-3.5 SR-5.4 | Equivalent wording used.
4.8 | The demographics collected will consist of the following points, the table below represents the name of the data, a validation if any and the requirement of the field to complete data entry: (see table in requirements document) | SR-3.4 | Equivalent wording used.
4.9 | The user can either cancel out of the process (Cancel) or they can save the information (Save). | SR-1.14 | Equivalent wording used.
4.1 | All data will be saved on the local system located at the client location. This includes all information contained on the patient intake form and the unique identifier generated based upon the patients name, date of birth and gender. | SR-1.5 | Equivalent wording used.
4.11 | A user can add to or update any of the existing patient’s information in requirement 4.8. | SR-3.5 | Equivalent wording used.
4.12 | The ClearView system must upload any new patients that are added to the system. | N/A | Business requirements dictate that this is no longer the direction we are going and when a centralized server-based solution is reconsidered the requirements will be modified comprehensively.
4.13 | The user should be able to go directly to the scanning functionality once a new patient has been saved. | SR-1.4 | Equivalent wording used.
SR-202 Rev. 004 Capture Functions Requirements |  | SR-206 rev. 000 New Requirement (Unless otherwise specified)
Reference | Description | New Reference | Justification
4.1 | The capture functionality should be available to any authenticated user of the ClearView system. | SR-2.7 | The access each role is granted will be included in the specifications.
4.1 | The capture search results should display the following data in columnar format: | N/A | Header statement for following requirements- shouldn't be ennumerated as a requirement.
4.11 | The default sort order of the list will be by date scanned in descending (most recent first) order. | SR-5.2 | Equivalent wording used.
4.12 | The search results list should be able to be sorted by any of the columns that are displayed in either ascending or descending order. | SR-5.1 | Equivalent wording used.
4.13 | Four functions should be able to be performed from the capture search results dialog: • Viewing the Report Pages • Viewing the Calibration Set Images • Viewing the Capture Images • Exporting of the subjects information including: • Finger Images • Calibration Images • Raw Calculated Data • Report Data • Calibration Data | SR-1.3 SR-1.4 | Equivalent wording used. Details of actual navigation steps will be documented in the specifications.
4.14 | Clicking the View Report button will load the report data for the selected patient and capture in the search results grid. | SR-6.4 | Equivalent wording used.
4.15 | Clicking the View Calibration button will display a dialog containing the Calibration images used during the selected scan. | SR-1.4 SR-4.11 | Equivalent wording used.
4.16 | Clicking the View Capture button will take the user to a dialog containing the images from the capture that was selected in the search results. | SR-1.4 SR-5.3 | Equivalent wording used.
4.17 | Clicking the Export Data button will present the user with a directory search dialog box which will allow them to specify where the export data will be saved for the selected subject. | SR-1.4 SR-5.11 SR-5.12 | Equivalent wording used.
4.18 | A new capture button should be available in the main menu of the application to initiate a new capture. | SR-1.4 | Equivalent wording used. Exact navigation functions will be documented in the specifications.
4.19 | A new capture can only be taken if the calibration routine has been successfully completed within the past 24 hours, and no more than 4 hours have elapsed since the last capture. If these are not true the calibration process must be run again. | SR-4.2 | Equivalent wording used. The specific interval that requires verifying the setup is a detail that will be included in the specifications.
4.2 | The capture must save the images in a 320 pixel wide by 240 pixel high Bitmap format image. | AR-1.3 | Documented in SR-207, equivalent wording.
4.2 | The new capture process must allow the user to select the patient from the system that the capture is being performed on. | SR-3.3 SR-5.8.1 | Equivalent wording used.
4.21 | The user can be located by searching on the following criteria or any combination thereof: • First Name • Last Name • Gender • Birth date | SR-3.3 | Equivalent wording used.
4.22 | The patient must first be searched on before a “New Patient” button will appear in order to prevent duplicate entry of patients. | SR-3.2 | The restriction of access to functionality in this case is a detail that will be included in the specifications. New requirement does identify the need for check to prevent duplication.
4.23 | The search results page should contain the following elements: • First Name • Last Name • Gender • Patient Id • Archived Flag | SR-3.3 | Equivalent wording used. The patient id is an internal database subject key that is not identifiable by the user and does not need to be displayed. Archiving requirements have not yet been determined and will be added to these requirements once established.
4.24 | The default sort on the new capture search results page will be first name. | SR-3.3 | Equivalent wording used. The default sort will be detailed in the specifications.
4.25 | The search results page should have buttons for three options available: • Create a new Patient • Get a new capture for the selected patient • View the selected patient's demographics | SR-1.4 SR-5.4 | Equivalent wording used. Details regarding the actual options displayed during navigation will be documented in the specifications.
4.26 | Clicking on the 'New Patient' button will take the user to the section of the application where patients are added to the system. | SR-3.1 | Equivalent wording used. The name of the button is a detail that will be included in the specifications.
4.27 | Clicking on the New Capture button will take the user to the new capture dialog for the selected individual. | SR-5.4 SR-5.8 | Equivalent wording used. The name of the button is a detail that will be included in the specifications.
4.28 | The new capture dialog should allow the display of twenty thumbnail images. These will consist of one image per finger for both hands, once for filtered and once for unfiltered. | SR-5.8.4 | Equivalent wording used.
4.29 | The dialog should display a live preview window which will show a live image of the finger before the energized image is captured. | SR-5.8.3 | Equivalent wording used.
4.3 | The dialog should show the captured bitmap image after the capture sequence has taken place. | SR-5.8.5 | Equivalent wording used.
4.3 | All images must be stored in the database, not on the file system. The file system can be used for temporary storage or calculations. | SR-5.7 | Equivalent wording used.
4.31 | The dialog should have information/messaging area on it so that messages can be displayed to the user as necessary during the capture process. | SR-5.8.4 | Equivalent wording used.
4.32 | The dialog should display a graphical representation of which finger is currently being scanned, as an orientation and verification point to the user. | SR-5.8.4 | Equivalent wording used.
4.33 | The currently selected image should have a red border around the thumbnail image. | N/A | The detail of the color of border around the finger image will be included in the specifications.
4.34 | The system should automatically calculate the angle of the image as well as the center point of the captured image. | SR-5.8.9.2 SR-5.8.9.3 | Equivalent wording used.
4.35 | The system should draw an oval around the calculated edge of the image. | SR-5.8.9.1 | Equivalent wording used (i.e., the shape actually drawn is an ellipse, not an oval).
4.36 | The system should display the angle of the image by drawing a line along the Y axis through the midpoint along the calculated angle. The line should start a few pixels above the drawn oval and extend below the drawn oval for a distance greater than that above the drawn oval. | SR-5.8.9.3 | Equivalent wording used.
4.37 | The system should draw a line along the X axis through the midpoint along the angle calculated + 90˚. The line will extend slightly beyond the drawn oval. | SR-5.8.9.2 SR-5.8.9.4 | Equivalent wording used.
4.38 | The oval and lines described above should be placed over the energized image as well as the finger image. | SR-5.8.8 | Equivalent wording used.
4.39 | Fields should be available on the dialog to edit the angel, the x coordinate and the y coordinate calculated by the system. The calculations must be able to be overridden. | SR-5.8.6 SR-5.8.7 | Equivalent wording used.
4.4 | All energized images, lit finger images, raw data and reported final data points are stored in the database. An audit trail will be kept for each datapoint stored in the database consisting of the following tracking points:  ID of the user that performed the image capture and initiated analysis.  Date and time of the capture/analysis down to the minute. | SR-5.8.2 SR-1.2 | Equivalent wording used.
4.4 | If the calculated values are overridden, the center point and angle generated by the user must be reflected in the crosshair axis and oval that is generated over the images. | SR-5.8.7 | Equivalent wording used.
4.41 | There should be an ability to turn on and off the oval and cross air on the images. | SR-5.8.10 | Equivalent wording used.
4.42 | The user should be able to select the image to capture by clicking on the thumbnail, this will make the clicked thumbnail the active capture point. | SR-5.8.11 | Equivalent wording used. The details of the process for recapturing an image will be included in the specifications.
4.43 | When a capture takes place, both the captured finger image as well as the captured energize image will display in the full size boxes. The energized image will also appear in the selected thumbnail. | SR-5.8.8 | Equivalent wording used. Details regarding the display layout will be included in the specifications.
4.44 | There should be a button to allow the user to return to the live preview in the case where the captured image must be re-captured. | SR-5.8.3 SR-5.8.11 | Equivalent wording used. The detail regarding the button to be displayed will be included in the specifications.
4.45 | The user can select a thumbnail at any time to display both the finger image and energized image that was captured. | SR-5.8.4 SR-5.8.11 | Equivalent wording used.
4.46 | Three action buttons will always be displayed for the user at the bottom of the dialog: | N/A | The detail of the displayed buttons and their function will be included in the specifications.
4.47 | An additional button 'Export Images' will be available only to EPIC system administrator users 'administrator'. This button will allow the user to export both the energized and finger images to the file system before an analysis is run. | SR-5.8.13 | Equivalent wording used.
4.48 | The 'administrator' user may also right click on a thumbnail to manually load an image from the file system. Right clicking on a thumbnail will cause an open file dialog to appear that will allow the user to select an energized image. No finger image will be loaded. | SR-5.8.12 | Equivalent wording used. The detailed process for loading images will be included in the specifications.
4.49 | The settings for the camera including the configurable values of voltage, brightness and gain should be displayed (Applies to version 2.0 and above). | SR-4.6 SR-4.7 | Equivalent wording used.
4.5 | A button should be available in the main menu of the application to search for a capture. | SR-1.3 SR-5.1 | Equivalent wording used. Requiring a "button" is a detail that will be included in the specifications.
4.5 | The ability to turn on and off the display of the centerpoint of the image and the angle of the image must be available. | SR-5.8.10 | Equivalent wording used.
4.51 | The system should make an attempt to reconnect to the scanner automatically if communication errors occur. | SR-1.7 SR-1.8 | Equivalent wording used. "Scanner" implies device firmware and device camera. The requirements have been separated and explicitly stated that it needs to maintain a single connection.
4.52 | The view capture dialog should allow the display of twenty thumbnail images. These will consist of one image per finger for both hands, once for filtered and once for unfiltered. | SR-5.3 | Equivalent wording used.
4.53 | The dialog will be loaded with the energized images from the data for the patient that was selected. Although there will be a placeholder for a large finger image, this will remain blank since finger images are not currently persisted. | SR-5.3 | Equivalent wording used. Remaining details will be implemented in the specifications.
4.54 | The system should display the saved angle as well as the midpoint values. | SR-5.3 | Equivalent wording used.
4.55 | There will be two buttons available to the user. The first button is a ‘Close’. This button will close the dialog and take the user to the main menu. The second button available to the 'Administrator' user only will be 'Resubmit'. Clicking this button will cause the capture selected to be resubmitted to the analysis process again. The result will not change the original records, but rather create a new set of data resulting from the resubmit process. | SR-1.14 SR-5.10.1 | Equivalent wording used. Requiring a "button" is a detail that will be included in the specifications.
4.56 | The ability to view the calibration will be available to any authenticated ClearView user. | SR-4.11 | Equivalent wording used.
4.57 | The dialog that appears will display thumbnail images of all ten calibration images that were used for the selected scan. | SR-5.3 | Equivalent wording used. Remaining details regarding display layout will be included in the specifications.
4.58 | A full size image (320 x 240) will be displayed in the middle of the dialog. | N/A | The detail regarding display sizes will be included in the specifications.
4.59 | Clicking on the individual thumbnail images will cause the image clicked to display in the full size image area, allowing the user to see a full size image of any of the thumbnails. | SR-4.11 | Equivalent wording used. Remaining details regarding display layout will be included in the specifications.
4.6 | The image that is currently displayed in the fill size image box in the middle of the dialog will have a thumbnail image bordered in red. | N/A | This is a detail that will be included in the specification.
4.6 | The ability to search for a capture should be available to any authenticated user of the ClearView system. | SR-2.7 SR-5.1 | Equivalent wording used.
4.61 | The patient name will appear at the top of the dialog. | SR-5.4 | Equivalent wording used.
4.62 | A single button will appear on this dialog, it will be labeled “Close” and will close the dialog and take the user back to the main menu. | SR-1.14 | Equivalent wording used.
4.7 | The user should be able to search for captures using the following criteria in any combination: • First Name • Last Name • Scan Start Date • Scan End Date | SR-5.1 | Equivalent wording used.
4.8 | The scan end date should default to no value. | N/A | Identifying the field format is a detail that will be move to the specifications.
4.9 | The scan start date should default to no value. | N/A | Identifying the field format is a detail that will be move to the specifications.
4.10.1 | Patient Name (First and Last) | SR-5.1 | Equivalent wording used.
4.10.2 | Treatment Id | N/A | Treatment ID will no longer be required to be displayed as a search result.
4.10.3 | Version of software the scan was run under | SR-5.1.3 | Equivalent wording used.
4.10.4 | Run Notes- When a scan is resubmitted, the Run Notes column should contain the verbiage: “Reprocessing of Treatment Id: XXXX” | SR-5.1.4 | Equivalent wording used.
4.10.5 | Scan Type | N/A | Scan type is no longer a requirement (i.e., treatment customization will not be pursued at this time).
4.10.6 | Scan Date- When a scan is resubmitted, the original date of the scan should be saved and displayed in the ‘Scan Date’ column. | SR-5.1.2 | Equivalent wording used.
4.10.7 | Archived Flag | N/A | Archiving requirements have not yet been determined and will be added to these requirements once established.
4.46.1 | A Close button – This will simply close the dialog. | SR-1.14 | Equivalent wording used.
4.46.2 | A Capture Image button – This button will begin the capture process for the selected image. | N/A | The detail of the displayed buttons and their function will be included in the specifications.
4.46.3 | A Submit for Analysis button – This button will submit the images collected to the analysis engine to perform the necessary calculations. The button will remain disabled until all 20 images have been successfully collected. | SR-5.8.15 | Equivalent wording used. The detail of the displayed buttons and their function will be included in the specifications.
SR-202-01 rev. 000 Scan Resubmit Requirements |  | SR-206 rev. 000 New Requirement (Unless otherwise specified)
Reference | Description | New Reference | Justification
Bullet 1 | A new button should be added to the view capture dialog to allow the image analysis to be rerun using the current algorithm implemented in the Analysis Functions. | SR-5.10.1 | Equivalent wording used. The details regarding buttons to be available will be included in the specifications.
Bullet 2 | The button will only appear for the administrator user. | N/A | Equivalent wording used. The details regarding buttons to be available will be included in the specifications.
Bullet 3 | If finger images are available, display them, otherwise show a blank window. | SR-5.10.3 SR-5.8.4 | Equivalent wording used.
Bullet 4 | Display angle and center point; Angles and center points can be changes and the modifications will be used in the new analysis. | SR-5.10.3 SR-5.8.8 | Equivalent wording used.
Bullet 5 | A new report and treatment Id should be created when the images are resubmitted through the system. | SR-5.10.2 | Equivalent wording used.
Bullet 6 | Save the version of software used to create the analysis results, this should be stored in the raw data table. | SR-5.10.3 SR-5.9.1 | Equivalent wording used.
Bullet 7 | Add version# to the capture search results page. | SR-5.1.3 | Equivalent wording used.
Bullet 8 | Remove both Capture ID and Patient Id from the search results. | N/A | Requirements should not delineate removal of previous requirements.
Bullet 9 | Change search dialog drop down to read “Search all scans” rather than “Search all treatment types”. | N/A | How search results are labeled will be included in the specifications.
Bullet 10 | Denote the new run as a re-run (possibly reference original data). | SR-5.10.2 | Equivalent wording used.
Bullet 11 | When the data is re-run it should be treated as an independent scan, a new treatment ID should be assigned and new results should be stored separate from the original results. | SR-5.10.2 SR-5.10.3 | Equivalent wording used.
SR-202-02 rev. 000 Persisting Finger Image Requirements |  | SR-206 rev. 000 New Requirement (Unless otherwise specified)
Reference | Description | New Reference | Justification
Bullet 1 | The saving of the image must work for all user roles (EPIC administrator, ClearView Administrator and ClearView User). | SR-5.8.2 | Equivalent wording used.
Bullet 2 | The image must be saved with the same detail level that is it is captured with. | SR-5.8.2 | Equivalent wording used.
Bullet 3 | The image must be retrievable during the viewing of a capture. | SR-5.3 | Equivalent wording used.
Bullet 4 | The View Capture window will be modified to display the finger image along with the energized image. | SR-5.3 | Equivalent wording used. Details regarding display layout will be included in the specifications.
Bullet 5 | The capture process window will be modified to display the finger image along with the energized image. | SR-5.8.5 | Equivalent wording used.
SR-202-03 rev. 000 Quick Export Requirements |  | SR-206 rev. 000 New Requirement (Unless otherwise specified)
Reference | Description | New Reference | Justification
Bullet 1 | This functionality should only be available to administrators. | SR-5.11 | Equivalent wording used.
Bullet 2 | Add a button to the Capture Search Results dialog that can be used to start an export of a subject's data. | N/A | The details regarding buttons to be available will be included in the specifications.
Bullet 3 | A button should appear for each search results row that appears in the Capture Search Results dialog.' | N/A | The details regarding buttons to be available will be included in the specifications.
Bullet 4 | The export process should allow the exporter to choose the destination location for the files to be saved. | SR-5.12 | Equivalent wording used.
Bullet 5 | The base folder that should be created for the export will follow this naming convention: {Subject Id}_Date and time of scan | N/A | The details regarding base folders for the export will be included in the specifications.
Bullet 6 | The process should export the following for the given scan that was selected. ◦ Calibration Images ◦ Energized Images (Filtered and Unfiltered) ◦ Finger Images (Filtered and Unfiltered, if available) ◦ The raw data for the calculation (in csv format) ◦ The report data | SR-5.11 | Equivalent wording used.
Bullet 7 | Within the folder selected, create the following folder directory structure: \CalibrationImages \FingerImages \EnergizedImages \RawData \ReportData | N/A | The details regarding folder directory will be included in the specifications.
Bullet 8 | Clicking the button should cause the following actions | N/A | The details regarding buttons to be available will be included in the specifications.
Bullet 9 | Display the output directory selection dialog | N/A | The details regarding output directory selection will be included in the specifications.
Bullet 10 | Export the data into the folder selected | N/A | The details regarding how export folders are selected will be included in the specifications.
Bullet 11 | Display a completion confirmation dialog. | N/A | The details regarding display messages for completion will be included in the specifications.
SR-204 rev. 006 (V. 2.1) Camera Functions Requirements |  | SR-206 rev. 000 New Requirement (Unless otherwise specified)
Reference | Description | New Reference | Justification
4.1 | The calibration process must energize the scanner 5 times before capturing the 10 images to use in the actual calibration calculation process. | SR-4.5 | Equivalent wording used.
4.1 | The process must prevent the user from exiting the process once it has begun until a final result has been reached. | SR-1.14 SR-4.1 | This requirement has changed because forcing the user to proceed through the entire setup verification process, when the user decides it was unintended, is no longer a requirement.
4.11 | The administrator and techadministrator users only, should be able to manually load calibration images, bypassing the auto calibration process. Once the user has added the images, a button to perform the calibration calculations should be available. | SR-4.15 | Equivalent wording used.
4.12 | The calibration process should not be accessible unless a scanner is attached to the system. If no scanner is found, a message should be displayed indicating such. | SR-1.7 | Equivalent wording used.
4.13 | The calibration dialog should allow the 10 captured images to be seen during the calibration process. | SR-4.9 | Equivalent wording used. Modified to remove the requirement to collect 10 images which will be included in the specifications.
4.14 | Calibration can be re-run at any time and for as many times as the user desires. Only the last successful calibration will be used for analysis purposes. | SR-4.1 SR-5.5 | Equivalent wording used.
4.15 | If a calibration process fails, a message should be displayed to the user. In addition the criteria that caused the failure should be placed in the alerts section of the application so they can be reviewed if necessary. The criteria would consist of the number of pixels that failed broken down by how many failed high and how many failed low. | SR-4.3 SR-4.13 | The portions of this requirement that were not included will be documented in the specifications.
4.16 | The calibration calculation is defined as follows: | N/A | Header statement for following requirements- shouldn't be ennumerated as a requirement.
4.17 | The calibration images and the calibration data should be able to be exported from the ClearView database. The export process will label the images with finger and hand references (1R). | SR-4.12 | Equivalent wording used. Details regarding image labeling will be included in the specifications.
4.18 | When a successful calibration is obtained, the images and a reference to the calibration is stored in the database. All scans performed after the calibration will reference the calibration. An audit trail will be kept for each image stored in the database consisting of the following tracking points:  ID of the user that performed the calibration.  Date and time of the calibration down to the minute. | SR-1.2 SR-4.10 SR-5.5 | Equivalent wording used.
4.19 | Upon a successful calibration, the images should be passed to the Matlab engine to generate the values necessary for the NS calculation. The NS calculation is specifically described in the Analysis Requirements document. | AR-3.14 | Documented in SR-207.
4.2 | The calibration must be calculated using 10 sample images captured sequentially. | SR-4.1 | The requirement to collect 10 images before analysis is being removed as this is a specification not a requirement. ClearView device testing will dictate the proper number of images to be collected for set up verification and the quantity of images will be documented in the specifications document.
4.2 | Each calibration image captured should be checked immediately for pass/fail status before grabbing the next image. If the image fails, the process should be immediately stopped. | N/A | The process flow for evaluating the images and for stopping/continuing the analysis will be documented in the specifications.
4.21 | The user should be able to select from the preconfigured calibration setting that were defined. | SR-4.8 | Equivalent wording used.
4.22 | Up to 4 preconfigured calibration settings should be available for use. The only users that can configure these settings are the administrator or techadministrator. The adjustable settings should be: 4.22.1 Brightness 4.22.2 Gain 4.22.3 Exposure Delay 4.22.4 Boost Voltage 4.22.5 The Default flag | SR-4.6 SR-4.7 SR-4.8 | Equivalent wording used. The requirement for the number of preconfigured settings will be included in the specifications.
4.23 | The user performing a calibration should be able to select from the preconfigured settings before starting the calibration process. | SR-4.8 | Equivalent wording used.
4.24 | The user should be able to re-run the calibration process in the case where it fails without having to exit and reload the calibration dialog page. | SR-4.1 SR-1.4 | The design will be implemented to ensure that user process flow is documented in the specifications.
4.3 | There should be a brief delay between image captures. | N/A | There is no requirement to delay between images. This requirement was originally implemented because the cause of a device performance problem had not been detected which has subsequently been resolved through hardware changes.
4.4 | The calibration process will be accessible to all users. | SR-2.7 | Details regarding user access will be included in the specifications.
4.5 | The button to access the calibration process must appear on the main menu. | SR-1.4 | Equivalent wording used. The specific buttons implemented will be included in the specifications.
4.6 | Calibration is required before a scan can be performed under the following circumstances: | N/A | Header statement for following requirements- shouldn't be ennumerated as a requirement.
4.7 | The date and time of a successful calibration as well as the images associated with the calibration must be stored in the database. No record of an unsuccessful calibration needs to be maintained. | SR-4.10 | Equivalent wording used.
4.8 | The calibration process should be controlled by a single button available to all users labeled 'Auto Calibration'. Clicking this button should do the following without user input: | SR-4.1 | Identifying a single button and the label of the button is a detail that will be move to the specifications.
4.9 | Present the user with a progress indicator of some sort so they are aware of where they are in the process. | SR-4.3 | Equivalent wording used.
4.16.1 | The calibration image will be broken down into four distinct regions of interest (the outer noise region, the inner noise region, the corona region and the high plateau region). | AR-2.2 | Documented in SR-207.
4.16.2 | Each of the areas is then compared to the threshold values for the respected area. | AR-2.4 | Documented in SR-207.
4.16.3 | If the total count of pixels that fail are equal to or greater than the limit specified for the given area, the image fails. | AR-2.4 | Documented in SR-207.
4.16.4 | The same process will be performed on all 10 images that have been collected. | SR-4.4 | Equivalent wording used. Modified to remove the requirement to collect 10 images which will be included in the specifications.
4.16.5 | If none of the 10 images fail, the calibration is considered a success. | SR-4.4 | Equivalent wording used. Modified to remove the requirement to collect 10 images which will be included in the specifications.
4.6.1 | If 24 hours have elapsed since the last calibration process has been performed. | SR-4.2.1 | This requirement has been modified to state "within reasonable intervals". The interval is a specification that will be dictated by device testing which will establish the appropriate interval. The appropriate interval will be documented in the specifications.
4.6.2 | If more than 4 hours have elapsed since the device was last accessed (aka. a scan was performed). | SR-4.2.1 | This requirement has been modified to state "within reasonable intervals". The need to re-establish appropriate setup verification is a specification that will be dictated by device testing which will establish the appropriate point for re-establishing the verification. The appropriate interval will be documented in the specifications.
4.6.3 | If the calibration settings have been changed in any way. | SR-4.2.2 | Equivalent wording used.
4.8.1 | Perform 5 warmup energize events. | SR-4.5 | Equivalent wording used.
4.8.2 | Collect and assess 10 images to use in the calibration process. | SR-4.1 | Modified to remove the requirement to collect 10 images which will be included in the specifications.
4.8.3 | Persist data if successful. | SR-4.10 | Equivalent wording used.
SR-205 Rev. 004 Application Functions Requirements |  | SR-206 rev. 000 New Requirement (Unless otherwise specified)
Reference | Description | New Reference | Justification
4.1 | The system should display the version number in the upper section of the main page. | SR-1.1 | Equivalent wording used. Actual display layout will be included in the specifications.
4.1 | The system must close after a specific amount of time of inactivity per HIPPA regulations. The initial amount of time should be 15 minutes. Inactivity is defined as no keyboard input or mouse clicks within the time frame. | SR-1.6.2 | Equivalent wording used. The detailed amount of time will be added to the specifications.
4.11 | The system should allow for notifications to the user to be displayed in a noticeable area regardless of the function that is currently being performed. | SR-1.1 SR-1.9 SR-1.11.2 SR-2.2.5 SR-2.3.3 SR-3.2 SR-3.6 SR-4.3 SR-5.8.4 SR-5.8.14 | Equivalent wording used. The notifications are now listed in the appropriate section of the requirements document instead of one comprehensive requirement as previously stated.
4.12 | The system should allow the device settings to be modified and persisted. These settings will be used by the application when controlling the device. This section is accessible to the "administrator" or “techadministrator” login only. These settings include: • Camera to use • Firmware Type • Brightness • Gain • Exposure Delay | SR-4.6 SR-4.7 | Equivalent wording used. Firmware Type refers to a historical design change (i.e., change to use C language) that will not be required for future software/hardware versions. This requirement is being deleted.
4.13 | A throttling mechanism should be put in place on the login screen to prevent a user from attempting to log in more than 3 times in a row with an invalid password. ‘In a row’ can be defined as within a 3 minute period. If three or more attempts are made with a single user id, the id should be logged and banned from logging in (even with the right password) for a period of 30 minutes. No error message should be shown to the user other than the standard invalid login message. The user id and date and time attempted should be logged in and login audit table and an error message should be presented to the user. If the user is banned there is no reason to log additional failures for that user. | SR-1.6.3 | Equivalent wording used. The details regarding what will happen when failed logins occur will be added to the specifications.
4.14 | The system must only be accessible to a user that has a valid User Id and password combination. | SR-1.6 SR-2.1 | Equivalent wording used.
4.15 | The system should allow the currently logged in user to be changed without requiring a shutdown of the application. | SR-2.8 | Equivalent wording used.
4.16 | Functions within the system must be delegated by role. There will be four roles in the system: • EPIC User- has access to scan functions only. • EPIC Administrator- has access to scan functions as well as user administrative functions like the Users functions. • System Administrator- has access to all areas of the ClearView software including Administrative functionality in the Capture, Camera and Report functions. • TechAdministrator- has access to camera setting functionality and will be shared with manufacturing parties and others that require access to camera settings. The TechAdministrator should not be able to see or change the core ‘Administrator’ user. | SR-2.5 | Equivalent wording used.
4.17 | The system should allow for the creation and management of system users. The following information should be captured for a user: (see table in document) | SR-2.2 | Equivalent wording used.
4.18 | The administrator and techadministrator logins should have the ability to capture an image from the settings page in order to perform testing of the device. | SR-1.10 | Equivalent wording used.
4.19 | The administrator and techadministrator logins shall have the ability to see the feedback from the firmware ‘Get’ command after an image capture has been performed on this page. This display would include the following values: • Device Status Code • LED Status • Exposure Duration • PWM0 Frequency • PON State • Boost Voltage Reading • PWM0 Width • Reset Condition Code • Firmware Version | SR-1.10 | Equivalent wording used. Details of firmware communications will be included in the specifications.
4.2 | The system must contain a main menu in order for the user to easily navigate through the application. | SR-1.3 | Equivalent wording used.
4.2 | The administrator and techadministrator logins shall have the ability to save the captured image (with or without the voltage and frequency displayed on the image). | SR-1.10 | Equivalent wording used. The requirement to record the voltage and frequency will be included in the specifications.
4.21 | The administrator and techadministrator logins shall have the ability to see a log of the commands that were sent to the device. | SR-1.10 | Equivalent wording used. Detail of firmware communications will be included in the specifications.
4.22 | The command log shall be able to be saved and cleared at any time. | SR-1.10 | Equivalent wording used.
4.23 | The ClearView system should require that a license file be present before the software can be used. | SR-1.11 | Equivalent wording used.
4.24 | The license file should be strongly encrypted as to not be able to be modified by the average user. | SR-1.11.3 | Equivalent wording used.
4.25 | The ClearView system should support three types of licenses, Basic, Full and Cardiovascular (CV). | SR-1.11 | Equivalent wording used. The type and labeling of licenses will be included in the specifications.
4.26 | A ClearView system utilizing a Basic license will display two output tabs upon completion of processing. One will contain the ClearView Basic printable report and the other will contain the working report (graphics, etc.). | SR-1.11.1 | Equivalent wording used. The functionality availble for each license will be included in the specifications.
4.27 | A ClearView system utilizing a Full or CV license will display two output tabs upon completion of processing. ClearView Users, ClearView Administrators, and TechAdministrators will have the Print Report and working report tables. EPIC Administrators will have the above mentioned tabs and a tab of admin pages as well as a tab of Biofields. | SR-1.11.1 | Equivalent wording used. The functionality availble for each license will be included in the specifications.
4.28 | The ClearView report displayed in the ClearView system utilizing a 'Full' license will be significantly different in format and content than the report under the 'Basic' license. The report formats will be developed based on Regulatory department feedback and documented in the final Application Functions Specifications. | SR-1.11.1 | Equivalent wording used. The functionality availble for each license will be included in the specifications.
4.29 | The ClearView report displayed in the ClearView system utilizing a 'CV’ license will be significantly different in format and content than the report under the 'Basic' license. | SR-1.11.1 | Equivalent wording used. The functionality availble for each license will be included in the specifications.
4.3 | The system must be Microsoft Windows based and be written using the .Net framework. | N/A | This detail will be included in the specifications.
4.3 | The license type being used (Full, Basic or CV) should be displayed to the user in an easy to recognize location. | SR-1.11.2 | Equivalent wording used. The type and label for each license will be included in the specifications.
4.31 | The ClearView system must interact with EPIC Central (the SaaS) application on a regular basis. If communication does not occur for a configurable amount of time or more, a scan cannot be performed. | N/A | The development of EPIC Central does not include documented requriements that meet EPIC's design controls. Once the work necessary to bring the DHF for EPIC Central is completed, any requirements that impact the ClearView Software development will be added to the ClearView Software requirements.
4.32 | The system must check prior to every scan if it is allowed to perform a scan. If it is not the scan must be prevented. | N/A | Billing requirements are not fully determined. Once they have been identified, these requirements will be updated and the necessary functionality developed under design controls.
4.33 | The system must allow for a configurable number of ‘Grace’ scans to be performed with the number of paid scans reaches 0. | N/A | Billing requirements are not fully determined. Once they have been identified, these requirements will be updated and the necessary functionality developed under design controls.
4.34 | The ClearView system must upload data to EPIC Central when a scan is completed, the data will consist of: (Patient demographic info, report data, raw data, and images. | N/A | The development of EPIC Central does not include documented requriements that meet EPIC's design controls. Once the work necessary to bring the DHF for EPIC Central is completed, any requirements that impact the ClearView Software development will be added to the ClearView Software requirements.
4.35 | Users with the EPIC Administrator role must be able to see the number of scans the device has available to use. | N/A | Billing requirements are not fully determined. Once they have been identified, these requirements will be updated and the necessary functionality developed under design controls.
4.36 | Users with the EPIC Administrator role must be able to access the EPIC Central UI from with ClearView to make purchases and manage their account. | N/A | The development of EPIC Central does not include documented requriements that meet EPIC's design controls. Once the work necessary to bring the DHF for EPIC Central is completed, any requirements that impact the ClearView Software development will be added to the ClearView Software requirements.
4.37 | ClearView must warn users when the number of scans available for a given device falls below a configured threshold. | N/A | Billing requirements are not fully determined. Once they have been identified, these requirements will be updated and the necessary functionality developed under design controls.
4.38 | Users with EPIC Administrator roles must be able to configure the scan count value that will trigger the “Low scan count” warning. | N/A | Billing requirements are not fully determined. Once they have been identified, these requirements will be updated and the necessary functionality developed under design controls.
4.4 | The system must persist data in a Microsoft SQL Server Database, version 2005 or greater. | N/A | This detail will be included in the specifications.
4.5 | The system should synchronize with an external controlling server. This server will be used to provide authentication for scans, archival functionality for images and other tasks. | N/A | Billing requirements are not fully determined. Once they have been identified, these requirements will be updated and the necessary functionality developed under design controls.
4.6 | The system should allow patients to be added and managed. The requirements for this functionality are defined in the Patient Services- Requirements document. | SR-3 | Equivalent wording used.
4.7 | The system should allow a calibration process to the performed on the scanning devices. The requirements for this functionality are defined in the Camera Functions- Requirements document. | SR-4 | Equivalent wording used. Calibration has been renamed setup verification.
4.8 | The system should allow for the capture of new images for a patient, as well as the analysis and results reporting for the images. The requirements for this functionality are defined in the Capture Functions- Requirements document. | SR-5 SR-6 | Equivalent wording used.
4.9 | The system should allow historical scans (i.e., patient energized images) and the analysis results (i.e., report displays) to be searched and reviewed by a user. The requirements for this functionality are defined in the Capture Functions- Requirements document. | SR-5.1 SR-5.3 SR-6.2 SR-6.3 SR-6.6 | Equivalent wording used.
SR-203 Rev. 004 Analysis Function Requirements |  | SR-207 Rev. 000 ClearView Software Requirements - Algorithm (unless otherwise specified)
Reference | Description | New Reference | Justification
4.1 | Image analysis must take place immediately upon request by the user after all of the images are captured. | SR-5.8.15 | Defined in SR-206, ClearView Software Requirements, equivalent wording used.
4.2 | Analysis requires a set of 10 filtered and 10 unfiltered images. | SR-5.8.4 | Defined in SR-206, ClearView Software Requirements, equivalent wording used.
4.3 | A valid calibration set must be current and available before analysis can begin. | SR-4.2 | Defined in SR-206, ClearView Software Requirements, equivalent wording used.
4.4 | The general data flow will be as follows: Collect images → Perform coefficients calculations → Store raw data (coefficient calculations) → Run raw data through analysis algorithm → Result in Response Scale Measurement (GSR measurements), NB score and LR score→ store analyzed data → produce a report. | N/A | This is a detail that will be included in the specifications.
4.5 | The system must calculate a noise value which will be referred to as the background noise level. This level should be the pixel intensity below which will be considered 0 or no intensity at all. This process is documented in the SR-203-01. | AR-3.1 AR-3.3 AR-3.5 | Reworded to just "determine an acceptable noise level".
4.6 | Energized finger images are to be divided for analysis into sectors; these sectors are to be unique for each finger and will be known as finger sectors. All calculations will be done based on the area within the specified sector. | AR-3.2 | Equivalent wording used. The definitions section provides the explanation for what a finger sector is.
4.7 | The number of pixels within a sector that are above the noise level should be measured, this coefficient will be known as the sector area. | AR-3.3 | Equivalent wording used.
4.8 | The area respective to the size of the sector must be calculated, this coefficient will be known as the normalized area. | AR-3.4 | Equivalent wording used.
4.9 | The average intensity of all pixels in a sector must be determined; this coefficient will be known as average intensity. | AR-3.5 | Equivalent wording used.
4.1 | The measurement of disorganization in a sector should be measured; this coefficient will be known as the entropy. | AR-3.6 | Reworded to include the detail that every angle is split into quarters and the mathematical functions are performed on quarters of angles for each angle in the sector.
4.11 | Measure the level of continuity along the radii of the sectors, this coefficient will be known as form. The form coefficient will encompass approximately 2/3 of the depth of the sector, a separate coefficient called Form2 will look at the outer 1/3 of the sector. | AR-3.7 AR-3.8 AR-3.10 | Equivalent wording used. "Continuity along the radii" is being phrased as "amount of change in pixel intensity..". Approximately 2/3 of the depth and 1/3 of the sector has been updated to "weighted and summed", a process which will be described in specifications.
4.12 | Breaks or lines in a sector will be known as Break Coefficient | AR-3.12 | Equivalent wording used.
4.13 | A measurement of the fractal dimension in a sector must be calculated, this coefficient will be known as fractal. | AR-3.13 | Equivalent wording used. Boxing is another word to depict fractal dimension.
4.14 | A measurement of the difference between a calibration image and the captured energized image per sector must be reported, this coefficient will be known as NS. | AR-3.14 | Equivalent wording used.
4.15 | A measurement of form in the first quadrant of the sector must be reported, this coefficient will be known as Form1_1. | AR-3.9 | Reworded AR-3.8 ties to AR-3.7 but breaks the calculation down into segments. This requirements just lists the segments where the new requirement only specifies that it must be done for each segment.
4.16 | A measurement of form in the first quadrant of the sector must be reported, this coefficient will be known as Form1_2. | AR-3.9 | Reworded AR-3.8 ties to AR-3.7 but breaks the calculation down into segments. This requirements just lists the segments where the new requirement only specifies that it must be done for each segment.
4.17 | A measurement of form in the first quadrant of the sector must be reported, this coefficient will be known as Form1_3. | AR-3.9 | Reworded AR-3.8 ties to AR-3.7 but breaks the calculation down into segments. This requirements just lists the segments where the new requirement only specifies that it must be done for each segment.
4.18 | A measurement of form in the first quadrant of the sector must be reported, this coefficient will be known as Form1_4. | AR-3.9 | Reworded AR-3.8 ties to AR-3.7 but breaks the calculation down into segments. This requirements just lists the segments where the new requirement only specifies that it must be done for each segment.
4.19 | A measurement of the average intensity in the first quadrant of the sector must be reported, this coefficient will be known as AI1. | AR-3.17 | Reworded uses the phrase "within each of the four concentric rings for each sector" rather than listing all four here. Sector concentric ring is defined in the definitions section.
4.2 | A measurement of the average intensity in the first quadrant of the sector must be reported, this coefficient will be known as AI2. | AR-3.17 | Reworded uses the phrase "within each of the four concentric rings for each sector" rather than listing all four here. Sector concentric ring is defined in the definitions section.
4.21 | A measurement of the average intensity in the first quadrant of the sector must be reported, this coefficient will be known as AI3. | AR-3.17 | Reworded uses the phrase "within each of the four concentric rings for each sector" rather than listing all four here. Sector concentric ring is defined in the definitions section.
4.22 | A measurement of the average intensity in the first quadrant of the sector must be reported, this coefficient will be known as AI4. | AR-3.17 | Reworded uses the phrase "within each of the four concentric rings for each sector" rather than listing all four here. Sector concentric ring is defined in the definitions section.
4.23 | A measurement of the number of pixels along radii that are within a specific intensity of each other must be calculated. In some cases this value may be zero. This coefficient will be known as Ring Thickness. | AR-3.15 | Equivalent wording used.
4.24 | A measurement of the number of the average intensity in the area that makes up the value for Ring Thickness must be reported, this will be known as Ring Intensity. | AR-3.16 | Equivalent wording used.
4.25 | The raw coefficients should be persisted in their original state in a table in the ClearView database. | SR-1.5 | Defined in SR-206, ClearView Software Requirements, equivalent wording used. The actual state of the values being stored will be included in the specifications.
4.26 | Z-scores will be calculated for each instance of all coefficients as described in the SR-203-01. | AR-4.1 | Equivalent wording used.
4.27 | A base score will be calculated for each finger sector, this score is a product of the weighting factor (which is unique to each coefficient), the calculated coefficient and the coefficient z-score as described in the SR-203-01. | AR-4.2 | Reworded to remove where the process is defined. The process will be clearly defined in the specification.
4.28 | Specific combinations of high scoring coefficients will be worth more than others. Rules will be put in place to enforce this. The rules are identified in the document SR-203-01. | AR-4.4 | Equivalent wording used. The rules will be included in the specifications.
4.29 | A scale will then be applied to the score that was created; the scale will reduce the results to a number between 0 and 25. The EPIC score scaling mechanism is documented in the SR-203-01. | AR-4.4 | The rules and the method for combining the total score to the rules (i.e., scale) will be detailed in the specifications. The EPIC score (the final score) and all associated documentation will be modified to be specifications documents.
4.3 | A new scoring process (Naive Bayse) will be added to the scoring calculation. This will be calculated completely independent of the EPIC scoring process. | AR-4.6 | Reworded to "odds ratio", naïve bayes is just one method of reaching the odds ratio and will be included in the specifications.
4.31 | The Naive Bayse (NB) score will be displayed in a new section in the ClearView report separated from the standard scoring section. This will be treated as an overall score. | N/A | Details of how the algorithm data is displayed to the user will be included in the specification.
4.32 | The Logistic Regression score will be calculated. The method of converting the EPIC score value to the final score for the report will be modified to use the process outlined in SR-203-01 to take population data into consideration. | AR-4.5 | Equivalent wording used. The "method of converting"will be included in the specification. See item 4.29 above.
4.33 | The method of converting the Logistic Regression score value to the final score for the report based on the process outlined in SR-203-01. This calculation will take the population data into consideration. | AR-4.7 | Equivalent wording used.
4.34 | A process for selecting which of the two scores for each measurement to select (EPIC or Logistic Regression) will be implemented into the score development algorithm. | AR-4.7 | The decision tree for determing the final score will be included in the specification.
4.35 | The ability to export data related to a scan must be implemented. Where applicable, the data must be exported in CSV format and files names will include finger and hand references (e.g., 1R) to the calibration images that are exported. The export functionality should only be available to the administrator user. | SR-5.11 | Defined in SR-206, ClearView Software Requirements. Reworded to exclude the detail of CSV format. The format will be included in the specifications.
4.36 | The following data should be exported: • The raw coefficients as well as the center point and angle of the images • The raw report data (data that would display in the report) • Calibration images • Energized images • Finger Images (if applicable) • Calibration data (the number of pixels that failed validation) | SR-5.11 | Defined in SR-206, ClearView Software Requirements. Equivalent wording used.
4.37 | Calibration images must be saved with the raw data. | SR-5.11.2 | Defined in SR-206, ClearView Software Requirements. Equivalent wording used.
4.38 | The final report data must be saved to the ClearView database. | SR-5.11.5 | Defined in SR-206, ClearView Software Requirements. Equivalent wording used.
4.39 | The raw Logistic Regression and EPIC scores and ranks must be able to be exported. | SR-5.11.4 | Defined in SR-206, ClearView Software Requirements. Equivalent wording used.
4.4 | An electronic record audit trail will be collected at the data collection points within the software. The audit trail will consist of who initiated the processing of the data (the current user ID) as well as the time and date stamp for the insertion of. The data collection points are: | SR-1.2 | Defined in SR-206, ClearView Software Requirements. Equivalent wording used.
4.41 | The user will have the ability to print a report of the Response Scale measurements.System should be able to graphically display the score results for organ systems, showing Physical and Autonomic scores and right and left values where appropriate. | SR-6.2 SR-6.5 | Defined in SR-206, ClearView Software Requirements. Equivalent wording used.
4.42 | The user will have the ability to view a graphical representation of the NS coefficient. | SR-6.8 | Defined in SR-206, ClearView Software Requirements. Equivalent wording used.
4.43 | The user will have the ability to view the energized image sectors in relation to a visual representation of the body known as the Biofield. | SR-6.9 | Defined in SR-206, ClearView Software Requirements. Equivalent wording used.
4.44 | The magnification algorithm used should implement a smooth scaling routine. | SR-6.9 | Defined in SR-206, ClearView Software Requirements.
4.45 | A report viewer will be used to allow the user to print and (based on user ID access) export the final ClearView Report. | SR-6.5 | Defined in SR-206, ClearView Software Requirements. The method of viewing the report will be included in the specifications.
4.46 | The ClearView report will be modified to have these sections in the "Full" license mode: | SR-6.1 | Defined in SR-206, ClearView Software Requirements. Specific license details have not been determined and this will be defined in the specification.
4.47 | The ClearView report will be modified to have these areas in the "CV" license mode: | SR-6.1 | Defined in SR-206, ClearView Software Requirements. Specific license details have not been determined and this will be defined in the specification.
4.48 | The ClearView report for the "Basic" licensed product will display only one section that contains Hand/Finger/Measurement # and the appropriate scores. | SR-6.1 | Defined in SR-206, ClearView Software Requirements. Specific license details have not been determined and this will be defined in the specification.
4.49 | The version of the firmware used in the device that was used to capture the patent images should be captured and stored with the treatment record. | SR-5.9.1 | Defined in SR-206, ClearView Software Requirements. Equivalent wording used. The firmware version is not specifically required as the software version is sufficient to understand how the image capture/analysis occurred.
4.5 | The version of device firmware as well as the version of ClearView software should be displayed when the results of a scan are viewed. | SR-6.4.2 | Defined in SR-206, ClearView Software Requirements. Equivalent wording used. The firmware version is not specifically required as the software version is sufficient to understand how the image capture/analysis occurred.
4.51 | The Id of the user that performed a calibration must be captured as well as the time and date of the calibration. | SR-1.2 | Defined in SR-206, ClearView Software Requirements. The ID of the user is unnecessary because the audit trail will capture at least the username and usernames are unique.
SR 203-01 rev. 000 Algorithm Requirements |  | N/A | This document was intended to be supplementary to the Analysis Requirements but is entirely details that will be included in the specification.