ORIGINATOR:

| Name | Job Title | Signature | Date |
| --- | --- | --- | --- |
| Brian Cullinan | Software Developer |  |  |

APPROVALS:

| Name | Job Title | Signature | Date |
| --- | --- | --- | --- |
| Nancy Rizzo | Founder, CEO |  |  |
| Lloyd Kurth | Technology Director |  |  |
| Dan Miller | Manufacturing Engineer |  |  |
| Jan Ayres | Project Implementation Manager |  |  |
| Andrea Miller | Quality Assurance Director |  |  |

Revision History

| Revision | Revision Date | Reason for Revision/Change Request | Revised By |
| --- | --- | --- | --- |
| 000 | 18-Sep-12 | Initial Release | L. Kurth B. Cullinan |

Table of Contents

1 Purpose	3

2 Scope	3

3 Definitions	3

4 System Architecture	3

4.1 Platform	4

4.2 Device Hardware	4

4.3 MatLab	4

4.4 SQL Server	4

5 System Design	5

5.1 ClearView application	5

5.2 Database and ORM	5

5.3 Third Party Libraries	6

5.4 Access Restrictions	6

5.5 Roles	6

5.6 Languages / Internationalization / Localization	6

6 Interface Design	7

6.1 Main Window	7

6.2 Settings	7

6.3 Users	7

6.4 Calibration	7

6.5 Patients	7

6.6 Capture	7

6.7 Reports	7

6.8 Logging & Diagnostics	7

6.8.1 Error handling	7

6.8.2 File logging	7

6.8.3 Auditing	7

7 References	7

# Purpose

# Scope

# Definitions

| Device Driver | A driver is a piece of software that runs on top of the Hardware Abstraction Layer within Windows. Windows provides and API for notifying the system when an event has occurred from a connected device. |
| --- | --- |
| Project | A project is a logical grouping of C# components and related files. A project is a component of a solution. |
| Solution | A solution defines all of the projects needed to produce the ClearView application. |
| Class / Object | A class or object is a logical grouping of functions. Each class defines its own level of exposure, public, private, internal, protected, and sealed. Depending on exposure, classes can be used by separate logical components of the ClearView solution. |
| App.config | A config file stores application level configuration options. These options can be encrypted, decrypted, and modified. The app.config should not store any sensitive information about the operation of ClearView, but is an easily accessible interface for changing the way some things work within the ClearView software. For example, the location of the log file that is generated while ClearView is operating. |

# System Architecture

ClearView interacts with a number of components with details described in other documents.  These components and their role in the ClearView software are described below, as well as necessary references listed in the references section of this document.

## Platform

ClearView is written in C# using Visual Studio and is compiled to run on Windows.  Full documentation for these products is listed in the references section.

ClearView will run properly on Windows XP / Windows 7 / and Windows 8, as the .Net Framework are supported on all of those platforms.

ClearView is heavily dependent on the .Net Framework version 3.5.  The .Net Framework is only support on the Windows operating system, therefore ClearView is currently restricted to Windows.

## Device Hardware

ClearView communicates with a microcontroller embedded within the ClearView device hardware.  This communication takes place over a Universal Serial Bus (USB) connection.  The outward facing microcontroller is an chip manufactured by FTDI.  Please see the device specification for more details on this hardware.

FTDI provides a C# “wrapper” around the device driver.  The wrapper has the ability to call methods within the driver software to perform functions such as “Write” and “Read”.  This writes data to a special section of memory that is then picked up by the driver software, converted in to a digital signal and sent to the device over a wire.  The FTDI chip then interprets that signal and performs commands programmed by the EPIC firmware.  For details on the firmware, please see the firmware specification.

Once the device performs a command, a response is generated in the same fashion in the reverse direction and is interpreted by ClearView.  All of the communication between device and ClearView is logged to a debug file.  ClearView uses some responses to acknowledge that the command was interpreted successfully by the firmware.

In general, when a command succeeds, the response is prefixed with “>,” followed by the command that was sent.  When a command fails, the response is prefixed with the characters: “?,”.  In the case of a failure, the entire “write/read buffer” (that special section of memory described above) should be cleared completely, so that the command can be reattempted later.

## MatLab

Please see the algorithm specifications for more details on the MatLab functions that occur throughout the imaging process.

## SQL Server

Structured Query Language (SQL) is a commonly used industry standard defined by ISO/IEC 9075:1992.  All data types used to store information throughout the software are SQL-92 compliant.

Microsoft’s own implementation of SQL-92 is called SQL Server.  ClearView is independent of what version it operates on because the SQL-92 standard is implemented in all versions, plus some additional sugar added and compatibility with later versions of the SQL language.

LLBLGen is an Object Relational Mapper (ORM) which tightly depends on what database server is used.  This relationship is described below.

# System Design

## ClearView application

The ClearView software ties all external components together and manages them with a single cohesive interface.  Each logical piece of the interface is described in detail below.

ClearView uses the .Net Framework to draw components to the screen and detect mouse and keyboard interactions with “buttons” and other interactive “objects”.

The .Net Framework handles all dynamic loading of libraries internally; ClearView performs no additional mechanism for loading components.

ClearView features a splash screen when it first starts; this indicates that libraries are being loading and the initial connection to the database is being attempted.

## Database and ORM

LLBLGen is an Object Relational Mapper (ORM).  It takes data from the database and converts it to object types that C# can understand.  It also takes data specified in C# that SQL can understand.  This process is described in detail below.

The following steps are detailed in procedural order in which they are used within ClearView code.

LLBLGen comes with a project designer.  This interface connects to various databases, but specifically SQL Server in the case of ClearView.

The layout of the database and all related design information is collected within the LLBLGen project designer.  This process maps SQL tables to C# objects, and SQL columns to C# properties.  As well as provides and interface to SQL stored procedures and any additional type mapping.

The ORM classes are generated automatically by LLBLGen, and the project is loaded within the ClearView solution.

In ClearView code, properties are accessed and set, when the ORM entity is saved the information is uploaded with an SQL command sent to the server.  These queries are parameterized for security reasons.

The queries sent and received through LLBLGen objects can be viewed using the flag specified in the app.config file.

<system.diagnostics>

<switches>

<add name="SqlServerDQE" value="3" />

## Third Party Libraries

ClearView uses a handful of third party libraries.  These libraries are referenced throughout the solution.

TVideoGrabber – Controls connecting to the camera.

This library has options for controlling brightness, exposure time, and frame rate.

Provides a standard interface between the camera and Windows forms.

This library also provides a control for displaying camera previews.

TVideoGrabber is used wherever the camera is displayed on the screen, the calibration tab, the settings tab, and the capture tab.

DirectShow – This library provides a standardized interface for connecting to cameras.

Infragistics – Renders controls to the user, such as the main menu Ribbon.

Crystal Reports – Displays the report.

## Access Restrictions

## Roles

## Languages / Internationalization / Localization

ClearView does not currently support internationalization.

# Interface Design

## Main Window

## Settings

## Users

## Calibration

## Patients

## Capture

## Reports

## Logging & Diagnostics

### Error handling

### File logging

### Auditing

# References

Appendix A – Requirements Traceability

Appendix B