ORIGINATOR:

| Name | Job Title | Signature | Date |
| --- | --- | --- | --- |
| Lloyd Kurth | Technology Director |  |  |

APPROVALS:

| Name | Job Title | Signature | Date |
| --- | --- | --- | --- |
| Nancy Rizzo | Founder, President and CTO |  |  |
| Brian Cullinan | Software Developer |  |  |
| Dan Miller | Manufacturing Engineer |  |  |
| Andrea Miller | Quality Assurance Director |  |  |

Revision History

| Revision | Revision Date | Reason for Revision/Change Request | Revised By |
| --- | --- | --- | --- |
| 000 |  | Initial Release | L. Kurth |

Table of Contents

1 Purpose	4

2 Scope	4

3 Definitions	4

3.1 References	4

3.2 Terms	4

4 ClearView Device	4

4.1 Device Firmware	4

4.2 Digital Camera	5

5 Host Environment	5

5.1 Windows	5

5.2 SQL Server	5

5.3 File System	5

5.4 .NET	6

5.5 FTDI Driver	6

5.6 FTDI C# Wrapper	6

5.7 Camera Driver	6

5.8 DirectShow	6

5.9 TVideoGrabber	6

6 Support Packages	6

6.1 Object/Relation Model (LLBLGen)	7

6.2 Crystal Reports	7

6.3 Printer Setup	7

6.4 PDF Creation	7

7 MATLAB	7

7.1 Setup Verification Analysis	8

7.2 Coefficient Analysis	8

8 ClearView Application	8

8.1 Users	8

8.2 Patients	8

8.3 Setup Verification	8

8.4 Scans	8

8.5 Reports	9

# Purpose

The purpose of this document is to describe the architecture…

# Scope

The scope of this document …

# Definitions

Throughout this specification document references are made to the following documents and terms.

## References

| Crystal Reports | The Crystal Reports .NET web site |
| --- | --- |
| CutePDF | The CutePDF web site |
| FTDI | The FTDI web site |
| LLBLGen Pro | The LLBLGen Pro web site |
| TVideoGrabber | Datastead Software’s TVideoGrabber web site |

## Terms

# ClearView Device

The ClearView Device is the proprietary hardware component of the ClearView system. It contains a glass electrode, a digital camera, microcontroller, USB communications, and power supply. The device is connected to the computer running the ClearView software via two separate USB connections. One is used to communicate with the onboard firmware running on the microcontroller. The other is used to control the camera. [Add references.]

## Device Firmware

Firmware executing on the microcontroller controls all device functions under direction of the ClearView application. The firmware implements the following primary functions:

Monitors boost voltage to insure it remains within safe operating range; if boost voltage is out of range, immediate terminates exposure and enters a fatal error state.

Starts an exposure and stops it after a configurable amount of time; exposure times are limited to a defined enumeration of durations.

Controls the frequency of the energy applied to the electrode; frequencies are limited to a defined enumerated set.

Controls the boost voltage of the energy fed to the high-voltage transformer which generates the energy applied to the electrode; boost voltage settings are limited to a defined enumerated set for normal operation.

Controls the on/off state of a flood LED which illuminates the surface area of the glass electrode.

Additional details of the firmware design can be found in SS-002, EPIC ClearView Firmware Design.

## Digital Camera

A digital video camera provides streaming video of images on the surface of the glass electrode. The camera is not controlled by the firmware. It is controlled entirely by the ClearView application through drivers provided by the camera manufacturer. To capture the still images used for image analysis, a frame is extracted from the video stream.

# Host Environment

The ClearView software runs on a standard Window-based personal computer. The host computer requires the usual user interface devices (keyboard, mouse, and monitor), two USB port to connect to the ClearView device, and a connection to a printer, which may be USB, network, or any other Windows-supported printer interface.

## Windows

The Microsoft Windows operating system provides the environment for the ClearView application. Windows 7, or later, will be installed.

## SQL Server

The standard Microsoft SQL Server relational database is used to store all operational data for the ClearView application. This includes configuration, users, patients, scanned images, setup verification images, and all calculated analysis data.

## File System

A standard Windows file system supports persistent storage of data. Persistent data includes the following:

Windows operating system files (compiled code, settings, temporary storage, etc.)

Compiled ClearView application files, loaded when the application starts

Configuration files required by the ClearView application

All files created and managed by the database server

Files created using ClearView export functions to archive application data

Installed supporting utilities

## .NET

The .NET Framework is Microsoft’s software environment that provides language independence, memory management, security and exception handling through the Common Language Runtime (CLR) application virtual machine (VM). The VM provides a high-level of abstraction through the .NET libraries allowing software developers to focus more on application logic rather than details of interfacing with the underlying operating system.

## FTDI Driver

The physical USB interface on the ClearView device that enables communication with the firmware is implemented with an FTDI chip. FTDI provide Windows drivers to enable software applications to interface with the chip. The ClearView system uses the FTDI D2XX driver.

## FTDI C# Wrapper

FTDI also provide a C# library that provides a managed .NET wrapper for interfacing to the D2XX drivers. The ClearView application utilizes this library.

## Camera Driver

The digital camera in the ClearView device connects to the computer via a USB interface. The software interface to the camera requires a Windows driver to communicate with the camera firmware. The camera driver is supplied by the camera manufacturer and installed in Windows. [Add specifics about current Phillips camera.]

## DirectShow

DirectShow is a Microsoft multimedia framework and API for working with media files and streams. In ClearView, DirectShow provides functionality to process the video stream from the camera. It connects with the camera through the camera driver.

## TVideoGrabber

TVideoGrabber is an SDK that supports video capture and recording within .NET. In ClearView, TVideoGrabber is used to capture frames of the video stream. The captured frames provide the still images which are analyzed to produce the patient scores. TVideoGrabber uses DirectShow to access the camera video stream.

# Support Packages

Several third-party support packages are used to supply off-the-shelf functionality that would otherwise need to be developed in-house.

## Object/Relation Model (LLBLGen)

LLBLGen Pro is an object/relational mapper and code generator that provides a simplified API for accessing the database. It allows the application to use objects that map to database tables rather than interfacing directly to the relational database through its primitive APIs. Data read from the database is marshaled into entity objects that represent the rows of a table. When an object is saved, the data is un-marshaled and written to the database, either creating a new row or updating an existing row.

LLBLGen Pro supports full reads of object hierarchies that represent the relationships among multiple tables. It also has full database transaction support and caching to improve performance.

LLBLGen Pro supports two types of modeling approaches: model first and database first. Using the model first approach, the software designer creates the application’s object model and the LLBLGen generates the database schema. Using the database first approach, the software designer creates the database schema and LLBLGen generates the object model. The model for ClearView is created with the latter approach while allows for more robust control over the database model.

## Crystal Reports

The Crystal Reports runtime engine for .NET component is used to generate the reports presented to the user. The reports are generated for viewing on-screen and can then be printed using standard print functions.

## Printer Setup

## PDF Creation

Portable Document Format (PDF) output is generated using CutePDF Writer, a freeware application that installs as a printer. Any printer output directed to the installed printer will result in the generation of a PDF file which the user can save to the file system.

# MATLAB

Mathworks MATLAB is a numerical computing environment and programming language. It allows for efficient development of algorithms for matrix management, linear algebra, and numerical computing. It is used within ClearView for image processing.

The full MATLAB development environment is used for developing the algorithms for image processing. But the development environment is not part of ClearView. A separate redistributable execution engine allows a MATLAB Compile Runtime (MCR) to be executed outside the full development environment. It is this execution engine that is included as a component of ClearView.

An MCR defines a Component Object Model (COM) interface that allows for invocation of its functions from other languages. The ClearView application code uses this COM interface to invoke the image processing algorithms that are essential to ClearView.

More details about MATLAB and its interface are provided in SS-004, ClearView Software Specification, Algorithm. That specification also provides details about the following two analysis components.

## Setup Verification Analysis

Setup Verification Analysis is the set of compiled MATLAB algorithms for analyzing setup verification images. During setup verification a set of images are captured from the ClearView device while a probe is resting on the glass electrode. Each of these images is processed through the MATLAB algorithms to determine if the ClearView device is properly configured. The algorithms process an image to obtain a set of characteristics (numerical data) that describe the image. These characteristics are then compared to a defined set of acceptable value ranges to determine if the image passes or fails. All images in the set of setup verification images must pass for the setup verification to pass.

After analyzing an image, the MATLAB component returns a pass-fail indication along with all the calculated characteristic values to the ClearView application. On user demand, the application displays all the calculated values so the user is able to make minor adjustments to the parameters that control the ClearView device. When a setup verification fails, the user can usually make adjustments so it will pass.

## Coefficient Analysis

Coefficient Analysis is the set of compiled MATLAB algorithms for analyzing energized finger images. After capturing the set of energized finger images that comprise a patient scan, each of the images is processed by the MATLAB algorithms to calculate a set of coefficients for the image. A set of coefficients is defined for each of the organs/structures. The calculated coefficients are returned to the ClearView application for additional analysis to determine final scores. These scores are then presented to the user.

# ClearView Application

## Users

## Patients

## Setup Verification

## Scans

## Reports